## What's new
What's new with IntelliVue…
This section lists the most important new features and improvements to the monitor and its user interface introduced with each release.
Software Release N.0
Software Release M.04
Software Release M.0
Software Release L.12
Software Release L.0
Software Release K.2
Software Release K.1
Software Release K.0
Software Release J.2
Software Release J.1
Software Release J.0
Software Release H.0
Software Release G.0
Software Release F.0
Software Release E.0
Software Release D.0
Software Release C.0
Software Release B.1
Software Release B.0
Software Release A.2
N.0/N.01
IntelliVue MX850 and MX750 The new high acuity stationary monitors are based on a similar hardware/software platform as the flexible care monitors MX400 - MX550. This platform offers a higher degree of cyber security and encryption capabilities. Moreover, the MX750/850 monitors offer a modernized user interface and experience to clinical users. IntelliVue AD75 and AD85 The new Active Displays are medical grade displays with integrated alarm lights and alarm speakers that are connected via a wired LAN connection to a host monitor MX750 or MX850. They provide an independent display to the host monitor. FMX-4 The new FMX-4 Module Rack is a next-generation 4-slot module rack with the capacity to support new types of measurement modules and to provide the bandwidth for high-density data transfer between an X3 and a host monitor. New Masimo O3 and Masimo CO2 Plug-in Modules The new 867184 O3 plug-in module, together with the Masimo O3 measurement device and sensors, provides a continuous measurement of the regional hemoglobin oxygen saturation of blood (rSO2) in the cerebral region under the sensors. The new 867185 CO2 plug-in module provides CO2 measurement using Masimo IRMA and Masimo ISA CO2 Analyzers. New Pressure and Temperature Labels The following new pressure labels are available for MX100 - 550, MX750/850 and X3 patient monitors:
| Label | Description
| IAP | Intra-abdominal pressure
| RFP | Right femoral arterial pressure
| LFP | Left femoral arterial pressure
| SCP | Spinal cord pressure
| UBP | Urinary bladder pressure
| Ppr | Pressure pre
| Ppo | Pressure post
| Pa1, Pa2, Pa3, Pa4 | Alternative non-specific arterial pressure labels
| Pv1, Pv2, Pv3, Pv4 | Alternative non-specific venous pressure labels
The following new temperature labels are available for MX100 - 550, MX750/850 and X3 patient monitors:
| Label | Description
| Tbladd | Bladder temperature
| Tfoley | Foley temperature
User Interface Enhancements With Release N.0, the Silence key has been renamed Acknowledge. Visual Alarm Indicators In the new default configuration, the background of the alarming measurement numeric flashes in the alarm severity color. Once the alarm is acknowledged, the background stops flashing, and the numeric flashes. Event Surveillance Enhancements • Event Surveillance now provides ten event groups instead of six. Six of these ten event groups can be activated. • Group activation can now be accessed directly in the Event Surveillance menu. • It is now possible to have more than one NER group. • The new group type CAR is now available.
back to top
M.04
New IntelliVue MMX Multi-Measurement Module The new IntelliVue MMX Multi-Measurement Module (867036) can simultaneously monitor ECG (using 3-, 5-, 6- or 10-electrode lead sets, including arrhythmia and ST monitoring), respiration, SpO2, NBP, two invasive pressures, temperature, and CO2. In combination with a host monitor, the MMX supports the current generation and previous generation Measurement Extensions.
back to top
M.0
New IntelliVue X3 MMS and Measurement Extensions The new IntelliVue X3 (867030) can simultaneously monitor ECG (using 3-, 5-, 6- or 10-lead sets, including arrhythmia and ST monitoring), respiration, SpO2, NBP, two invasive pressures, temperature, and CO2. The multi-touch screen allows easy interaction by sliding and tapping with one or two fingers. The new measurement extensions offer additional measurements and can be powered from the X3 battery during transport. • 867039 Hemodynamic Extension - cardiac output, temperature, two pressures • 867040 Capnography Extension - mainstream/sidestream CO2, cardiac output, temperature, two pressures • 867041 Microstream Extension - Microstream CO2, cardiac output, temperature, two pressures Alarm Advisor The alarm advisor functionality provides feedback on recurring and continuous alarm limit violations, for a specific measurement over a period of time. The information provided supports you in adapting alarm limits more specifically for individual patients. Electronic Strip Recordings with On-Screen Review Electronic strip recording allows alarm-triggered and manually started electronic strips to be captured in the monitor print database. They can then be reviewed on the monitor, and printed in the form of reports when a printer is available. Source for QRS Tone Configurable In previous releases, the source for the QRS tone automatically followed the alarm source. It can now also be set manually to HR or Pulse.
back to top
L.12 (MP5/5SC and XG50 only)
Temporal Temperature Measurement The choice for intermittent or spot check temperature measurements is extended with the Exergen TemporalScanner™ thermometer. The thermometer measures the patient's temperature on the forehead, over the temporal artery, using infrared technology. The result is displayed on the display of the thermometer and is transmitted to the monitor. XG50 With the new XG50 patient monitor you can collect vital signs and save them as SpotCheck records. In addition, you can use Guardian Early Warning Scoring to calculate a score based on the vital signs.
back to top
L.0
Compact Gas Analyzer Plug-in Module The new IntelliVue G7m Gas Analyzer module is a compact, double-width, plug-in module that measures five anesthetic agents, N2O, and the respiratory gases oxygen (O2) and carbon dioxide (CO2). Masimo Rainbow SET® The choice of different pulse oximetry technologies is extended with Masimo Rainbow SET. In addition to oxygen saturation, pulse rate and perfusion index, Masimo Rainbow SET's multi-LED technology provides measurement of: • Total Hemoglobin • Oxygen content • Carboxyhemoglobin • Methemoglobin • Pleth variability index • Acoustic respiration The availability of additional measurements depends on the options purchased. Note that the instructions for using Masimo Rainbow SET are provided in an addendum to the Instructions for Use. EEG Enhancements For the new M1027B module, the IntelliVue software L.0 provides two new functionalities: • aEEG - The aEEG presentation is a trend display of the amplitude-integrated EEG. The aEEG presentation is optimized for two and four channels. A trend of the sum of the electrode impedances of the respective lead is shown below the aEEG presentation as a quality indicator that supports interpretation of the aEEG. The monitor stores 24 hours of aEEG and electrode impedances for all 4 channels. • Burst Suppression Ratio (BSR) - The burst suppression ratio indicates the amount of time within an interval spent in the suppressed state. The BSR numeric presents the ratio as a percentage. A BSR of 100% corresponds to an isoelectric (flatline) EEG. Protocol Watch Sepsis Screening Enhancements A new protocol is now available that supports the Surviving Sepsis Campaign 2012 guidelines. Remote Alarm Device The new remote alarm device provides additional audio and visual indications of alarms remote from the bedside. When alarms are audible at the remote alarm device, they can be muted at the monitor, to provide a low-noise patient environment. When the caregiver goes to the bedside they can switch audible alarms back on again with a single touch selection. RFID Barcode Reader Support In situations where a barcode could previously be read into the monitor, for example for patient data at admission, RFID data can now also be read. A new barcode reader is available that incorporates RFID reading technology and can be used for traditional barcodes and RFID tags. ECG Enhancements - Extended Selection for Va and Vb Chest Leads For 6-lead placement you now have an extended selection of chest leads: V1 to V9 and V3R to V6R. - Specific Label Assignment for V Lead With standard 5-lead placement, the generic V lead can be configured to display with its specific position (V1 to V9 and V3R to V6R). - STE Map The STE Map adds gender-specific STE limits to ST Map. ST values violating these limits are indicated in red. - 12-Lead Order Management A pop-up key provides access to the order window from the Information Center. You can assign an order number from the pending list to a 12-Lead ECG waiting to be stored. This functionality requires corresponding configuration and licenses at the Information Center. Check with your Philips representative for availability. NMT Enhancements - Automatic Measurement Start The monitor can now be configured to automatically start a measurement cycle immediately after a successful calibration. - Easy Manual TOF Measurement During Measurement Cycle During a measurement cycle in Auto mode you now have a Single TOF key that allows a manual measurement at any time with a single key action. - Additional TOF Repeat Time The repeat time for TOF measurements can now also be set to 2 minutes. - Improved Access to Repeat Time Repeat Time has been moved forward in the setup menu to give you easy access for changes during a measurement cycle. Lead Selection for Respiration Measurement It is now possible to choose between two different lead vectors for the impedance respiration measurement. Stat Admit Function The new Stat Admit function allows you to admit a patient with a temporary patient identification. It can be used in cases when the patient ID is unknown or when the data is not yet available.
back to top
K.2
Cableless Respiration Measurement The new CL Respiration Pod can be used with the monitor. It is a battery-powered, cableless, respiration rate measuring device that is applied to the patient's chest using an adhesive attachment. In addition to the respiration rate it can also provide posture information and optionally a pulse measurement. Hexad Derived 12-lead ECG The Hexad ECG lead system uses a 6-lead set connected to four limb electrodes (using Mason-Likar placement) and two chest electrodes. The algorithm derives the four remaining V-leads to provide a non-diagnostic 12-lead view, including ECG waves and ST measurements. Accelerated Measurement for NBP An accelerated measurement is now available. This allows a faster measurement result but requires that the patient keeps the limb in question still. The accelerated measurement must be switched on in Configuration Mode. New EEG Module The new EEG module (M1027B) provides enhanced functionality with four channels of realtime EEG waves. The monitor automatically detects whether a two or four channel module or trunk cable is connected. Remote Applications for MX450/500/550 Applications hosted on a customer-supplied Citrix XenApp® server or the PIIC iX Web can be displayed and operated on the patient monitor screen. IntelliVue XDS Solution The IntelliVue XDS Solution is now also available for the MX400-550 models allowing you to view an independent monitor screen on an external display. Depending on the configuration you can also operate the monitor from the external display. Tympanic Temperature You can now measure tympanic temperature with the new tympanic thermometer. The thermometer provides measurement values on its built-in display and communicates them to the monitor via a connected cable. The result of this measurement can be automatically adjusted to correspond to a different body reference site. The alarm limits can be set at the monitor.
back to top
K.1 (MX400 - 550 only)
MX500/MX550 In this release the MX500/MX550 patient monitors join the IntelliVue monitor family. The monitors are optimized for transport within the hospital environment. The MX500 has a 12-inch display and the MX550 has a 15-inch display. Both models have three integrated plug-in module slots for up to three measurement modules, or the IntelliBridge interface module (865115), or recorder module (M1116B/C). The MX500/550 have the option to be operated with a single, user-replaceable battery. EC10 IntelliBridge Interface Board The functionality of the IntelliBridge EC10 plug-in module - interfacing data from a connected external device into the monitor - is now available in an embedded board. The connector for the optional interface board is accessed at the rear of the monitor. Wireless Remote Control The remote control, previously only available with a USB cable connection, can now also be used wirelessly, with a short-range radio connection, as with the models MX600-800.
back to top
K.0 (MX400/450 only)
MX400/MX450 In this release the MX400/MX450 patient monitors join the IntelliVue monitor family. The monitors are optimized for transport within the hospital environment. The MX400 has a 9-inch display and the MX450 has a 12-inch display. Both monitors can optionally be equipped with an integrated recorder and have the option to be operated with a single user-replaceable battery.
back to top
J.2 (MP5/5SC/5T only)
Tympanic Temperature You can now measure tympanic temperature with the new tympanic thermometer. The thermometer provides measurement values on its built-in display and communicates them to the monitor via a connected cable. The result of this measurement can be automatically adjusted to correspond to a different body reference site. The alarm limits can be set at the monitor. Cableless Respiration Measurement The new CL Respiration Pod can be used with the MP5/MP5SC/MP5T monitors. It is a batterypowered, cableless, respiration rate measuring device that is applied to the patient's chest using an adhesive attachment. In addition to the respiration rate it can also provide posture information and optionally a pulse measurement. EWS and SpotCheck Enhancements - General Enhancements • The new Cableless Respiration Pod can be used in the EWS procedure. • Revised and additional SpotCheck Trend views are available to support various workflows. • When the monitor is used with GuardianSoftware, notifications for the currently selected patient can be received from GuardianSoftware and viewed on the monitor. • Notifications can be sent from the monitor to GuardianSoftware. - User Awareness • When the monitor is used with IntelliVue GuardianSoftware, user identification and authentication functionality is available that allows users to log-in at the monitor and defines permissions for various actions at the monitor. • When user identification/authentication is in use, users with the appropriate permission can validate data. • Permissions can also be assigned to allow users to make changes to SpotCheck data, to acknowledge notifications from GuardianSoftware and to send notifications to GuardianSoftware. - Patient Context Awareness • Before confirming the admission of a patient in Quick Admit a patient search is made in the GuardianSoftware database or the monitor's multi-patient database. • SpotCheck records, EWS protocol selection and patient data changes can be synchronized between GuardianSoftware and the monitor. SpotCheck records for the current patient that were made on other monitors or at GuardianSoftware can be viewed on the monitor. • Hospital-configurable SmartKeys, "Escalate-Notification" and "Acknowledge-Notification", support hospital workflow quality management and enable paging from the patient monitoring device. - EWS Configuration Enhancements • Automatic saving of a SpotCheck record at the end of the freeze time or when a new patient is selected is now configurable. • A variety of waves can now be displayed in the EWS entry screen. • The range of numerics and enumerations that are supported has been increased. • The lifetime of a measurement can now be configured independently of the configured freeze time. Accelerated Measurement for NBP An accelerated measurement is now available. This allows a faster measurement result but requires that the patient keeps the limb in question still. The accelerated measurement must be switched on in Configuration Mode.
back to top
J.1
Switch Port Mapping The Link Layer Discovery Protocol (LLDP) is a vendor-neutral protocol used by network devices which enables the exchange of information between neighbors in a wired network (enables Switch Port Mapping). The LLDP mode displays its status together with additional information in the Network Status window acquired by the switch to which the monitor is connected when using LAN. New Software options for MX600/700 • #A08 – Eight Real-time Wave segments (may not be commercially available in all geographies; check availability for your country) • #C07 – Advanced Event Surveillance
back to top
J.0
Flexible, Patient-oriented Workflow (with PIIC iX only) With the IntelliVue Information Center iX, more flexible, patient-oriented workflows are supported for admitting, tracking and transferring patients. This allows you to adapt monitor use to the patient data flow models in your specific facility and department. Managing Patients When the patient monitor is connected to an IntelliVue Information Center iX, an extended comparison of patient data allows more intelligent admission and transfer procedures. A Find Patient key can be used to search through the patient lists at the Information Center and other connected systems and take over the corresponding data found. Flexible Transfer A transfer can be initiated in the unit at the monitor, on patient arrival or patient departure, or centrally at the Information Center. Managing Equipment When used together with an Information Center iX, extended equipment management functionality is now available. Depending on how equipment is used in your facility, there are various ways to associate devices with patients, beds or monitors. Additionally there are mechanisms to automatically free up equipment that is no longer used. Using the new Equipment window you have an overview of all equipment being used for a patient, together with patient and caregiver data. New NMT Module A new NMT module enables the evaluation of muscle relaxation of patients under Neuromuscular Block by measuring the strength of muscle reaction after electrically stimulating the dedicated motor nerve. Alarm Functionality Enhancements • Pause Yellow Alarms (with PIIC iX only) – up until now, the Pause Alarms / Alarms Off key switched off or paused all red and yellow alarms. Now an alarm priority configuration setting is available to have only yellow alarms affected. • Pause Alarms / Alarms Off not allowed – when the alarm priority is set to Not Allowed, alarms can no longer be switched off or paused. The Pause Alarms / Alarms Off key is disabled. • Pause Alarms / Alarms Off key can be removed from the screen – a configuration setting removes the permanent key from the screen to avoid unintentional switching off of alarms. Alarms can still be switched off or paused, in the Alarms menu, under Main Setup. • Alarms paused after Standby mode – alarms are automatically paused for 1 minute after Standby mode to allow time for patient connection. Capture 12-Lead ECG Enhancements (with PIIC iX only) The existing 12-Lead ECG capture functionality has been extended to offer: • Download of analysis results from the IntelliVue Information Center. • Download of 12-Lead captures from the IntelliVue Information Center for review at the patient monitor. • Analysis results included in reports and additional report selections with ST Map. • Remote operation of the 12-Lead Export function at the IntelliVue Information Center. • Remote operation of the 12-Lead Lock/Unlock function at the IntelliVue Information Center. • New filter settings that are used as a default for future 12-Lead captures. New "End Afib" and "End Irregular HR" Alarms The end of an atrial fibrillation or irregular HR phase is announced with an **End AFIB or **End Irregular HR alarm. The delay time before the alarm is announced can be configured. Information Line Changes Patient Category is now displayed as a symbol in front of the patient name and cannot be changed from the Information line. Paced Mode is no longer displayed in the Information line; it appears next to the HR numeric. Depending on your configuration, you will see the current Profile name, and/or the Screen name in the Information line. The Profiles menu has been simplified in Monitoring and Configuration Mode. Sepsis Protocol Configurability The SSC Sepsis protocol can now be configured, to customize it for the specific treatment measures used in your facility. This can include changes to limits for measured or manually entered values, changes to the Severe Sepsis Screening criteria, and changes to the recommendations in the Sepsis Resuscitation Bundle and the Sepsis Management Bundle. Integrated Pulmonary Index for Microstream CO2 An Integrated Pulmonary Index (IPI) numeric is provided, that is an indication of the patient's overall ventilatory status based on four measurement parameters: etCO2, awRR, pulse rate and SpO2. Cardiac Output (transpulmonary method) - Changes to Calculation Basis for Extravascular Lung Water Index and Preload Indexes EVLWI is now calculated using Ideal Body Weight (IBW). ITBVI and GEDVI are calculated using Ideal Body Surface Area (IBSA). IBW and IBSA are based on the patient category, gender and height that are entered for the patient. Improved Visibility of Gridlines in Overlapping Waves Visibility of the gridlines for overlapping waves has been improved, and you can now configure gridlines to be shown in white, if you prefer, in a thin or thick style. Timeout for "Aged" Numerics NBP and SpO2 numerics from intermittent measurements can be configured to be grayed out or to disappear from the screen after a set time. This avoids older numerics being misinterpreted as current data. The time can be set individually for SpO2 and NBP in Configuration mode.
back to top
H.0
MX800 new Intellivue Patient Monitor New Design for the User Interface The user interface for the patient monitors has been redesigned to bring the presented information into the foreground, letting the structural elements such as keys and frames retreat into the background. Additionally special regard was given to improved usability and compatibility with standard software products. New tcGas Solution The IntelliVue TcG10 is a new device using state-of-the-art technology for the measurement of transcutaneous O2 and CO2 partial pressure in neonates, pediatrics and adults. The measurement values, alarms and messages are displayed at the patient monitor, using the IntelliBridge EC10 module as the interface. 12 Lead ECG Enhancements You can now preview the 12-Lead ECG at the bedside before you send it to the Information Center. A new dedicated 12-Lead ECG Report can be printed in diagnostic quality from the bedside. New Remote Control Device The remote control provides you with direct access to five hard keys, a navigation knob and a numeric keypad which can also be used for alphanumeric entry. The hard keys include Silence, Alarms Off/Pause Alarms, Main Screen and a SmartKeys key that displays a block of SmartKeys specially configured for remote tasks. Vital Signs Trend Based on NBP Measurement Times You can now have the NBP measurement determine the interval between the columns on the Vital Signs trend. Each NBP measurement will generate a column in the trend table. The values for the other measurements are added to the column to provide a complete vital signs set for the NBP measurement time. Dynamic Wave Sizing In addition to fixed size waves, Screens can now contain a Dynamic Wave area where the waves will automatically adapt their size to the space available. For example, a wave will increase in size when one or more of the configured waves is not available. BIS Impedance Values in Sensor Window In addition to symbols indicating the impedance status of the electrodes, the impedance value is also displayed in the BIS window.
back to top
G.0
Short Range Radio Interface for MP5 and X2 A telemetry transceiver with a short range radio adapter can now be assigned to an MP5 or X2 which is connected (via Companion Mode) to a larger host monitor (MP20 - MP90). The data from the telemetry transceiver are displayed on the host monitor with a minimal delay. IntelliBridge Module (MP40/50, MP60/70, MP80/90) The new IntelliBridge EC10 module transmits information from a connected external device to your monitor. Waveforms, measurement numerics, settings, and alarms can be displayed on the monitor, passed on to an Information Center, and included in trends. Intravascular Oxygen Saturation Measurement (MP40/50, MP60/70, MP80/90) The new M1011A SO2 module extends the continuous measurement of intravascular oxygen saturation to the MP40/50 MP60/70, MP80/90 class of IntelliVue Patient Monitors. The measurement label is selectable for different measurement locations and is provided to the SSC Sepsis protocol. Histogram Trend Trended values can also be viewed in the form of a histogram. The SpO2 histograms can be trend histograms or real-time histograms with 1 second samples. Car Seat Assessment Record In Event Surveillance, in the NER group, you can run a Car Seat Assessment Record (CAR). This is a special period of event surveillance for neonates during a car seat test. During the CAR period a realtime SpO2 histogram is also generated with 1 second samples. IntelliVue XDS Solution Using the IntelliVue XDS solution it is possible to view an independent monitor screen on an external display. The XDS solution consists of a medical grade PC-based hardware platform, XDS application software and the XDS connectivity option on the monitor. Depending on the configuration you can also operate the monitor from the external display. The XDS must be connected to the same Local Area Network (LAN) as the monitor. It is also possible to use an existing PC, connected to the same LAN, to host the XDS Application software. New Fields in the Admission Form A Middle Name field is now available in the admission form - whether it appears is configurable. Two additional ID fields, Lifetime Id and Encounter Id, can also be configured to appear and their names can be customized to fit hospital requirements. SpO2 Enhancements • Additional SpO2 Labels - SpO2 pre and SpO2 post labels have been added. • New manual measurement mode when telemetry devices are connected via a short range radio link. • Signal Quality Indicator displayed with the SpO2 numerics. ECG Enhancements • New */**Afib yellow alarm when an atrial fibrillation waveform is detected. • QT View window - shows current wave and baseline wave with Q and T points marked so that you can verify that the QT algorithm detects correct Q and T points. USB Interface A new USB interface allows use of USB printers, bar code readers and input devices such as a keyboard or mouse.
back to top
F.0
X2 Multi-measurement Module (M3002A)
You can connect the new X2 to an IntelliVue patient monitor, where it acts as a Multi-Measurement Module (MMS), providing measurements, trends and patient information to the MP20/30, MP40/50
and MP60/70/80/90. You can also use the X2 as a stand-alone, highly portable monitor with a display and alert capabilities, running on a rechargeable battery or mains power from the optional external
power supply (M8023A). With these two use modes the X2 provides an ideal transport solution - changing its function from MMS to transport monitor and back to MMS as required.
Name Changes for MMS, FMS, etc.
In the process of introducing the X2 the naming for the Multi-measurement Servers and associated equipment has been changed.
The following table shows the new names and the abbreviations used - the abbreviations have been kept as they were to ease the change:
| Previous Name |  | New Name | abbreviations | Product Number
| Multi Measurement Server |  | Multi measurement Module | MMS | M3001A X1
M3001A X2
| Multi Measurement Extension |  | MMS Extension | - | M3012A/4A/5A/6A
| Flexible Module Server |  | Flexible Module Rack | FMS | M8048A
| Measurement Server Link |  | Measurement Link | MSL | -
Protocol Watch
The SSC Sepsis protocol is now available on all IntelliVue Monitors (MP20 to MP90). Depending on the monitor model, and in some cases on the option purchased, you may have only the
screening phase implemented or the complete protocol. If a patient has been screened on arrival on a monitor with the screening option, then transferred to Intensive Care
when Sepsis is confirmed, the current state, the settings and the log from the SSC Sepsis Protocol will be transferred with the MMS to the new monitor.
Should it be necessary, you can also select the correct phase of the protocol at the new monitor manually.
QT/QTc Interval Monitoring
QT interval monitoring is used to assist in the detection of prolonged QT interval syndrome. The QT interval is measured and the heart-rate corrected QT value, QTc, is calculated.
The monitor also provides a ΔQTc value which tracks variation in the QT interval in relation to a baseline value. High alarm limits can be set for QTc and ΔQTc enabling alarm notification.
Manual Data Entry
Measurement values (for example lab data or manually measured temperatures) can be entered manually into the monitor and stored in the database. They can then be used for trends, reports and
also be displayed as a numeric on screen, if required.
Horizon Trend Enhancements
The trend indicator arrow indicates how the patient trend has developed in a set time period. This period is now configurable and can be set to 10 minutes, 5 minutes or 2 minutes.
The horizon (or baseline) can now be set to a range or to a specific value.
Connecting an MP5 to a Monitor (Companion Mode is indicated)
The MP5 monitor can be connected to one of the MP20 to MP90 monitors (then called the host monitor) and used like an MMS. This provides monitoring continuity in transport situations allowing
measurement data and patient demographics from the MP5 to be transferred to another monitor. When the MP5 is connected to a host monitor, no alarms will be announced on the MP5.
Pairing X2 and MP5 with a Monitor
The X2 and MP5 (with ITS) can be declared as telemetry devices at the Information Center and paired with a monitor. The data from the MP5/X2 will then be shown on the same sector on the
Information Center as that from the monitor.
NBP - Programmable Measurement Sequence
Up to four measurement cycles can be set up which will run consecutively. For each cycle you can set the number of measurements and the interval between them. By setting the last cycle to run
continuously you can have regular measurements continue after the sequence has run.
New Printing Options For Reports
Reports can now also be printed via an external PC-based print server or to an internal print database, when no printer is available (for example, during transport).
As soon as a printer becomes available, the reports stored in the database will print automatically.
Cardiac Output - Automatic Detection Of Injections
An Auto mode is now available in the cardiac output measurement which allows you to make multiple cardiac output measurements without having to press the Start button every time.
The injections are automatically detected after the first measurement of a series has been started.
Manual Pairing At Monitor
Previously, devices could only be manually paired at the Information Center. You can now pair devices at the monitor as long as the monitor is already connected to the Information Center.
Moving Windows And Menus
Windows and menus can now be moved on the monitor screen. Using touch or a mouse you can select the title of a window and then “drag” it across the screen.
Some positions on the screen are not allowed, such as ones overlapping the alarm field or the monitor info line.
Invasive Pressure Changes
Wave cursor - A cursor is now available on the realtime pressure wave to allow you to define a position and store the corresponding value.
Reduced alarm limit steps in lower ranges - The alarm limits setting in 2 mmHg steps is now possible in ranges up to 50 mmHg (previously up to 30 mmHg).
New Patient Check
The monitor can be configured to ask you in certain situations: after a specified power-off period, after a specified standby period and when no basic vitals (HR, RR, Pulse, SpO2, NBP) have been
measured for a specified period, whether a new patient is now being monitored. By selecting the Yes pop-up key you can discharge the previous patient and reset settings.
New Smart Keys
A SmartKey is available for the new Lab Data entry function and for New Lead Setup when a new ECG lead set with fewer leads than previously is being used.
Better Visibility Of Gridlines
The brightness of the gridlines on the realtime waves has been increased for better visibility.
back to top
E.0
Protocol Watch - SSC Sepsis Protocol
ProtocolWatch is a clinical decision support tool. It allows you to run a clinical protocol which can monitor developments in the patient’s condition.
The SSC Sepsis Protocol runs on the ProtocolWatch application and is used in screening for severe sepsis and monitoring its treatment.
Telemetry Data Enhancements
When a telemetry device is paired with the monitor, you can now change settings such as heart rate alarms or relearn arrhythmia for the telemetry device at the monitor.
When the ECG source changes certain relevant settings will be synchronized between the monitor and the telemetry device to provide settings continuity.
Telemetry alarms and INOPS will also be indicated on the monitor, in addition to the main indication at the Information Center.
Pulse From Noninvasive Blood Pressure
In the process of making the NBP measurement, a pulse value can be derived and displayed. The pulse value is displayed together with the time the measurement was made.
Alarms Enhancements
• In addition to the standard blue INOPs, some INOPs can now be configured as red or yellow INOPs to provide a severity indication (ECG leads Off, NBP Cuff overpress, NBP Cuff Not
Deflated). A small number of INOPS are always yellow or red to indicate a severity corresponding to red and yellow alarms.
• For the invasive pressure measurement, the extreme pressure alarms Extreme High and Extreme Low can be made available for your monitor in Configuration Mode and are additional to the
standard High and Low limit alarms.
• The Review Alarms window now shows when the monitor was switched on (after being switched off for longer than 1 minute) and any changes made to the Standby and paired status.
• In order to improve alarming on asystole under certain conditions, you can set Asystole Detect. in Configuration Mode to Enhanced. In enhanced mode an asystole alarm will be
suppressed for up to five seconds if a valid beat-to-beat Pulse is detected from a Pressure.
SpO2 Perfusion Change Indicator
The perfusion change indicator is a graphic symbol which shows the change in the perfusion value, relative to a reference value which you can set.
Pressure Zero Stored in Module
The Zero will be stored in the pressure module when it is unplugged and moved to another monitor, as long as the transducer is still connected to the module when it is plugged in again.
Respiratory Waves Can Be Displayed Filled
For certain respiratory waves (CO2, AWF, AWP, and AWV) the presentation of the wave trace can be changed in configuration mode to display filled waves.
BIS Module in MP60/70
The BIS module can now be used in the integrated module slots in the MP60 and MP70.
IntelliVue Instrument Telemetry Wireless Network Available Outside USA
Built-in adapters (for MP20/30) and external adapters (for MP40/50) are available for the new IntelliVue Instrument Telemetry network for use outside the USA.
ECG
• New ECG Filter Setting The new Extended Monitoring filter setting is for use on pediatric and neonatal patients when diagnostic quality is required but low frequency interference or a wandering
baseline may be expected. The upper edge frequency is the same as the Diagnostic setting and the lower edge frequency is the same as the Monitoring setting.
• Pacer Spikes are shown in a different color to the ECG wave. They are shown in white unless the ECG wave is white, then they are shown in green.
• 6-lead cable is now supported.
PiCCO Enhancements
• New measurement values (not available in the USA): GEF (Global Ejection Fraction), PVPI (Pulmonary Vascular Permeability Index)
• New continuous measurement: PPV (Pulse Pressure Variation)
• Right-Left Shunt detection (not available in the USA)
• GEDV, GEDVI, EVLW and EVLWI now also available in the USA
• SVV, PPV and dPmax can now also be measured prior to a CCO calibration. (dPmax not available in the USA.)
New SmartKeys
Two new SmartKeys allow you to unpair a telemetry device from the monitor and determine whether the Information Center will continue to receive data from the monitor or the telemetry device.
Lead Placement selection - EASI or Standard - can be accessed with a SmartKey. There are now separate SmartKeys for Starting or Stopping an NBP measurement and for Starting a STAT measurement.
Two new SmartKeys provide shortcuts to the Record Waves and Record Vitals functionality.
back to top
D.0
M1014A Spirometry Module
The spirometry module produces a real time wave for flow, volume and pressure of respiratory gases together with numerics for analysis of ventilatory mechanics.
It is designed to be used in combination with Philips-branded airway flow sensors and combined CO2/airway flow sensors.
M3014A Capnography Extension
The M3014A Capnography Extension offers an additional measurement method - sidestream CO2 with the M2741A sensor.
M8016A D80 Intelligent Display
The D80 Intelligent Display can be used as a third main display with the MP90 monitor. You then have three displays able to be configured individually and to be operated independently.
IntelliVue 802.11 Bedside Adapter
The 802.11 Bedside Adapter is a wireless ethernet adapter which can be built in to the patient monitor (option J35). It provides wireless connectivity to standard IEEE 802.11a/b/g wireless networks.
Improved Multiple Display Support
• Wide and Tall screen layouts are supported which allow screen content to be spread over two screens next to each other (wide screen) or one above the other (tall screen).
• An additional main display used for a surgeon can be configured not to show alarm information. The permanent keys and SmartKeys can also be configured not to display.
• Input devices can be assigned to an operator, independent of physical connector location.
New SmartKeys
For the new spirometry module there is a SmartKey to provide direct access to the spirometry data window. For microstream and sidestream CO2, sampling can now be suppressed directly with the
SmartKey CO2 Pump Off.
back to top
C.0
IntelliVue MP80
The MP80 patient monitor is a new addition to the IntelliVue patient monitor family. The functionality is similar to that of the MP70 but with the flexibility of component modularity as in the MP90.
M3014A Capnography Extension
The M3014A Capnography Extension offers a new measurement combination - mainstream CO2, cardiac output, invasive pressure and invasive pressure/temperature.
M8045A Docking Station
The docking station provides quick mounting and connections for the MP20/MP30/MP40/MP50 in a one-step operation. By placing the monitor on the docking station and closing the lever you can
make the connection to AC power and to a network, if present.
BIS Interface Board
This interface board allows use of Bispectral Index monitoring with the MP20 and MP30.
Timers Application
The new Timers application allows you to set timers to notify you when a specific time period has expired. The timers can have varying characteristics
and can be located on the Main screen for easy viewing.
IntelliVue Instrument Telemetry IIT (USA only)
Wireless network capabilities via the IntelliVue Instrument Telemetry network using a built-in interface (MP20/MP30) or an external adapter (MP40/MP50).
Remote Applications
Remote applications can now be embedded in a monitoring Screen. The display colors for remote applications have been enhanced.
Basic Operation
• Selection and order of SmartKeys can now be configured in the monitor’s Configuration mode.
• A new SmartKey exits Standby mode at the gas monitor.
• Direct operation of a second independent display with the MP90, using standard input devices.
• Networked monitors can now show Telemetry information for the “Own Bed” in an overview window or embedded in a Screen.
• Overview bed information can be configured to display in the colors used at the Information Center (with Information Center System G or higher).
• Support for Unit-based Care Group model for up to 64 beds (with Information Center System G or higher).
• Battery status information on the main screen now includes estimated monitoring time available during battery charging.
• Barcode reader support during patient admission.
• Direct entry to graphical trends when a limit alarm is selected in alarm review and to the event episode window when an event alarm is selected.
• MP90: screen trends can now also be viewed on the second display.
Trends
• In graphical trends, a segment menu allows direct adjustment of trend scales, automatic scale optimization, expanded view for an individual segment.
• In graphical and horizon Screen Trends a cursor is available.
• In the Vital Signs window, the values can be shown with the parameter color.
Events
• Events can be configured to be signaled as alarms (advanced event surveillance only).
• Deviation triggers can be set which trigger an event when measured values change by a defined amount over a set time (advanced event surveillance only).
• Up to six event groups can be active simultaneously (advanced event surveillance only).
Measurements
• Pulse Pressure Variation is a new derived measurement calculated from beat-to-beat arterial pressure values.
• MAC value calculations
Applications
• Improved Drug Calculator meeting JCAHO requirements.
• The ST Map application shows ST changes over time in two multi-axis spider diagrams.
• Freeze and measure waves on the Main Screen.
back to top
B.1
IntelliVue MP20/MP30 :
The MP20/MP30 patient monitor is a new addition to the IntelliVue patient monitor family. It is smaller and lighter than the MP40/MP50 monitors and can be powered by battery. It can be used with
the measurement server and server extensions and has a built-in recorder.
Basic Operation
• New Measurement Selection window makes it easier to resolve measurement label conflicts
• New Previous/Next Screen function provides access to the ten most recently modified Screens
• Wave speeds can now be set for individual wave channels
Trends
• Cursor in graphical trend window improves navigation in the trends database
• Vital Signs and Graphical Trend screen elements can be embedded on a Screen
• New band style format for displaying trends of measurements with multiple numerics
• Horizon trend is a new format for screen trends, showing the deviation from a stored baseline
• Aperiodic measurements now stored with a timestamp in Vital Signs
• New symbol representing NBP measurements in Graphical Trends
Measurements
• ST numerics in the Alarm Limits window can be shown and hidden
• ST Point can be set directly by selecting a numeric value
• New NBP countdown timer shows the time remaining until the next NBP measurement in a series
• Networked monitors can now show Other Bed information embedded on Screen
• Additional labels available for Pressure and Temp
• Additional options for SpO2, support for an extended list of accessories
• Suppress Zero function for Microstream CO2
• M1020B support in internal Slot of MP60/70
Applications
• Drug Calculator can now be configured to include a list of commonly-used drugs
• Cursor in the Loops window improves navigation through the stored loops
• Volume-flow loops added to the list of respiratory loops types
• Loops Report for documentation of stored loops
• EEG CSA can now be called up and viewed in a window over the currently displayed Screen
back to top
B.0
IntelliVue MP40/MP50
The MP40/MP50 patient monitor is a new addition to the IntelliVue patient monitor family. It uses the same measurement devices as the MP60/MP70/MP90 monitors
and shares the same technological platform and user interface, but is more compact in size and can be operated by battery.
M3012A Measurement Server Extension
The new Hemodynamic Measurement Server Extension extends measurement capability by adding two additional pressures and Cardiac Output.
M1020B SpO2 Module
New SpO2 measurement module, M1020B, enables dual SpO2 measurement without the need to use the VueLink module. Two options are available:
• Option A01 for use with Philips reusable and disposable sensors and Nellcor “R-Cal” disposable sensors.
• Option A02 for use with Nellcor OxiMax sensors, including the MAX-FAST forehead sensor.
• M1020B Option A02 for use with Nellcor OxiMax sensors may not be available in all countries.
PV Loops
The PV Loops application compares graphic representations of airway waves to help detect changes in the patient airway condition.
High-resolution waves per Screen
The number of high-resolution waves that can be shown on a Screen is increased, limited only by the Axx Option purchased.
Alarms symbols
New alarm symbols are introduced, and “short” yellow alarms were renamed “onestar” yellow alarms (yellow arrhythmia alarms).
Aperiodic measurements available as Screen Trends
Patient trend information for NBP, C.O., C.I., and Wedge can now be permanently displayed on the Screen in tabular and graphical form.
back to top
A.2
12-Lead ECG recordings
12-Lead ECG waves and numerics can be sent to a connected recorder
High-Resolution Trend Report
high-resolution trend report can be sent to a connected printer ST Snippets ST snippets, showing a one second wave segment for each measured ST lead, can be
permanently displayed on the Screen or called up as required.
EEG Wave Speed
New EEG-specific wave speeds have been added to the list of wave speeds available
Drug Calculator
This new feature helps you to calculate drug dosages for your patients. On-Screen Calculator: a mathematics calculator can be used on the Screen
Visitor Screen
This new Screen is designed to hide sensitive patient information from the Screen. Monitoring and alarm generation function as usual.
Touch selection volume control
The volume of the audio prompt given when a screen element is selected is now adjustable
VueLink interface
The VueLink on-screen appearance and controls are improved
M3001A
Trend upload from the Multi-Measurement Server (M3001A) improved
Screen Trends
Lets you display patient trend information in graphic form permanently on the Screen
Alarm Limits Page
Lets you view and control alarm settings for all measurements in one window
New Option for Event Surveillance
A new neonatal event review option #C04 is introduced
Second display
To simultaneously show two different Screens, a second display can be connected to the MP90. The second display is for viewing only.
M3046 Support
The M3001A is supported with the new M3046A SW E.00.23
back to top

## SW IntelliVue
Intellivue SW/HW Revisions History
|  |  | M3001A | M3002A | 867036 | 867030/33 | 866060/62 | 866061 | 866064/66 | 866470/71 | 867130/33 | 866468 | 865240 | 865241 | 865242 | 865243 | M8048A | M8105A | M8105AT | M8105AS | M8001A/2A | M8003A/4A | M8005A/7A | M8008A | M8010A | M8016A | Comment
|  |  | MMS | M8102A | MMX | X3/MX100 | MX400/450 | MX430 | MX500/550 | MX850/750 | AD75/85 | FMX-4 | MX800 | MX700 | MX600 | FMS-4 | FMS-8 | MP5 | MP5T | MP5SC | MP20/30 | MP40/50 | MP60/70 | MP80 | MP90 | D80 | Related FCO/SB
|  |  | (HW rev see | X2/MP2
comment)
Software A.x is no longer supported | A.0 | NLS Rev. 6.08 | A.05.51 | x | x | x | x | x | x | x | x | x | x | x | x | x | A.05.50 | x | x | x | x | x | A.04.20 | x | x | x | FCO86200088
|  |  | A.05.52 | [13 empty] | (A.00.01) | [5 empty] | A.05.50 |  |  |  | FCO86200089
|  |  | A.05.53 | [19 empty] | (A.00.01)
| 2002-08-01 |  | (HW A)
| A.0 |  | A.05.52 | x | x | x | x | x | x | x | x | x | x | x | x | x |  | x | x | x | x | x | A.06.07 | x | A.06.07 | x | FCO86200101
|  |  | A.05.53 | [13 empty] | A.05.50 | [5 empty] | (A.00.01) |  |  |  | FCO86200102
| 2002-10-01 |  | (HW A) | [13 empty] | (A.00.01) | [9 empty] | FCO86200108
| A.1 |  | A.10.14 | x | x | x | x | x | x | x | x | x | x | x | x | x | A.10.14 | x | x | x | x | x | A.10.15 | x | A.10.15 | x | FCO86200153
|  |  | A.10.16 | [13 empty] | A.10.16 | [5 empty] | (A.00.01) |  | (A.00.01) |  | FCO86200152
| 2003-01-01 |  | (HW A) | [13 empty] | (A.00.01)
| A.2 | NLS Rev. 6.10 | A.20.27 | x | x | x | x | x | x | x | x | x | x | x | x | x | A.20.27 | x | x | x | x | x | A.20.46 | x | A.20.46 | x | FCO86200267
|  |  | (HW A) | [13 empty] | (A.00.01) | [5 empty] | (A.00.01/2/3) |  | (A.00.01/2/3) |  | FCO86200287
2003-07-01
| A.2 |  | A.20.27 | x | x | x | x | x | x | x | x | x | x | x | x | x | A.20.27 | x | x | x | x | x | A.25.13 | x | A.25.13 | x | FCO86200344
| with 100Mhz only |  | (HW A) | [13 empty] | (A.00.01) | [5 empty] | (B.00.04/5/6) |  | (B.00.04/5/6)
2003-11-01
| B.0 | NLS REV 6.20 | B.05.64 | x | x | x | x | x | x | x | x | x | x | x | x | x | B.05.64 | x | x | x | x | x | B.00.02 | x | B.00.02 | x | FCO86200344
| with 50Mhz only |  | B.05.73 | [13 empty] | (A.00.01) | [5 empty] | (A.00.01/2/3) |  | (A.00.01/2/3)
| 2003-11-01 |  | (HW A)
| B.0 |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | x | x | B.00.05 | x | - | x | SB86200084
| with 50Mhz only | [21 empty] | (A.00.01/2/3) |  |  |  | FCO86200476
2004-09-01
| B.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | x | x | B.00.06 | x | - | x | support ELO Touch
| with 50Mhz only | [21 empty] | (A.00.01/2/3)
2006-05-06
| B.0 |  | B.05.64 | x | x | x | x | x | x | x | x | x | x | x | x | x | B.05.64 | x | x | x | x | B.05.67 | B.05.67 | x | B.05.67 | x | FCO86200344
| with 100Mhz only |  | B.05.73 | [13 empty] | (A.00.01) |  |  |  |  | (A.00.03) | B.00.04 |  | (B.00.04/05/06) |  | FCO86200319
| 2003-11-01 |  | (HW A) | [23 empty] | FCO86200476
| B.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | x | B.05.71 | B.05.71 | x | B.05.71 | x | SB86200084
| with 100Mhz only | [20 empty] | (A.00.03/04/05/08) | (B.00.04/05/06/18/22/23) |  | (B.00.04/05/06/18/22/23)
2004-09-01
| B.0 Patch ELO |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | x | - | B.05.72 | x | B.05.72 | x
| with 100Mhz only | [21 empty] | (B.00.04/05/06/18/22/23) |  | (B.00.04/05/06/18/22/23)
2009-06-01
| B.0 Patch ELO |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | x | B.05.73 | - | x | - | x | SB86201021A
| for MP40/50 | [20 empty] | (A.00.03/04/05/08/10/12)
2009-09-01
| B.1 | NLS Rev. 6.25 | B.10.71 | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | B.10.71 | - | - | x | - | x | SB86200045
|  |  | (HW A/B) | [17 empty] | (A.00.04)
2004-06-01
| B.1 |  | B.10.81 | x | x | x | x | x | x | x | x | x | x | x | x | x | B.10.80 | x | x | x | B.10.79 | B.10.80 | B.10.82 | x | B.10.82 | x | SB86200043
|  |  | (HW A/B) | [13 empty] | (A.00.01) |  |  |  | (A.00.04) | (A.00.03) | B.00.04 |  | (B.00.04/05/06) |  | FCO86200446
SB86200128
| 2004-07-01 | [25 empty] | SB86200043
| B.1 |  | B.10.83 | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | - | B.10.84 | B.10.84 | x | B.10.84 | x | SB86200084
|  |  | (HW A/B) | [18 empty] | (A.00.03/04/05/08/10/12) | (B.00.04/05/06) |  | (B.00.04/05/06/18/22/23) |  | FCO86200514
2004-09-01
| B.1 |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | B.10.85 | B.10.85 | B.10.85 | x | B.10.85 | x | SB86200159
[21 empty] | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) |  | (B.00.04/05/06/18/22/23)
2004-11-01
| B.1 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | B.10.86 | B.10.86 | B.10.86 | x | B.10.86 | x | FCO86200644
[21 empty] | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) |  | (B.00.04/05/06/18/22/23) |  | FCO86200645
2006-02-01
| B.1 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | - | B.10.87 | - | x | - | x | SB86201021A
(A.00.03/04/05/08/10/12)
2009-09-01
| C.0 | NLS Rev. 6.30 | C.00.81 | x | x | x | x | x | x | x | x | x | x | x | x | x | C.00.81 | x | x | x | C.00.82 | C.00.82 | C.00.82 | C.00.82 | C.00.82 | x | SB86200218
|  |  | (HW A/B) | [13 empty] | (A.00.01) |  |  |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) |  | SB86200242
SB86200249
| 2005-06-01 | [25 empty] | FCO86200601
| C.0 |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | C.00.89 | C.00.89 | C.00.89 | C.00.89 | C.00.89 | x | SB86200270
[21 empty] | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2005-10-05
| C.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | C.00.90 | C.00.90 | C.00.90 | C.00.90 | C.00.90 | x | FCO86200644
[21 empty] | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) |  | FCO86200640
2006-02-01
| C.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | - | C.00.91 | - | - | - | x | SB86201021A
(A.00.03/04/05/08/10/12)
2009-09-01
| D.0 | NLS Rev. 6.35 | D.00.54 | x | x | x | x | x | x | x | x | x | x | x | x | x | D.00.57 | x | x | x | D.00.58 | D.00.58 | D.00.58 | D.00.58 | D.00.58 | D.00.58 | SB86200337
|  |  | (HW A/B/C) | [13 empty] | (A.00.01) |  |  |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | SB86200370
2006-02-01
| D.0 Patch |  | D.00.61 | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | - | - | - | - | - | - | FCO86200673
|  |  | (HW A/B/C) | [23 empty] | *)Mandatory upgrade to D.00.61
2006-09-01
| D.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | D.00.59 | D.00.59 | D.00.59 | D.00.59 | D.00.59 | D.00.59 | FCO86200774
[21 empty] | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | FCO86200787
| 2007-04-01 | [25 empty] | SB86200605
| D.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | x | x | x | - | D.00.60 | - | - | - | - | SB86201021A
(A.00.03/04/05/08/10/12)
2009-10-01
| E.0 | NLS Rev. 6.40 | E.01.24 | x | x | x | x | x | x | x | x | x | x | x | x | x | E.01.22 | E.01.26 | E.01.26 | x | E.01.24 | E.01.26 | E.01.24 | E.01.24 | E.01.24 | E.01.24 | SB86200509
|  |  | E.01.28 | [13 empty] | E.01.28 | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | E.01.26 | E.01.28 | E.01.28 | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | FCO86200743
| 2006-12-01 |  | (HW A/B/C) | [13 empty] | (A.00.01) |  |  |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) |  |  |  | SB86200551
| E.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | E.01.30 | E.01.30 | x | - | - | - | - | - | - | SB86200606
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16)
April-07
| E.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | E.01.32 | E.01.32 | x | E.01.31 | E.01.31 | E.01.31 | E.01.31 | E.01.31 | E.01.31 | FCO86200860
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2008-02-01
| E.0 Patch |  | - | x | x | x | x | x | x | x | x | x | x | x | x | x | - | - | - | x | - | E.01.32 | - | - | - | - | SB86201021A
(A.00.03/04/05/08)
2009-10-01
| F.0 | NLS Rev. 6.45 | F.01.24 | F.01.33 | x | x | x | x | x | x | x | x | x | x | x | x | F.01.24 | F.01.34 | F.01.34 | x | F.01.33 | F.01.33 | F.01.35 | F.01.35 | F.01.35 | F.01.35 | SB86200662
|  |  | (HW A/B/C) | F.01.36 | [12 empty] | (A.00.01) | F.01.37 | F.01.37 |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
| 2007-09-01 |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16)
| F.0 Patch |  | - | F.01.38 | x | x | x | x | x | x | x | x | x | x | x | x | - | F.01.39 | F.01.39 | x | F.01.40 | F.01.40 | F.01.40 | F.01.40 | F.01.40 | F.01.40 | minor changes introduced
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | with new ST
2007-10-01
| F.0 Patch |  | F.01.41 | F.01.43 | x | x | x | x | x | x | x | x | x | x | x | x | - | F.01.43 | F.01.43 | x | F.01.43 | F.01.43 | F.01.42 | F.01.42 | F.01.42 | F.01.42 | FCO86200859
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2008-02-01
| F.0 Patch |  | F.01.41 | F.01.43 | x | x | x | x | x | x | x | x | x | x | x | x | - | F.01.43 | F.01.43 | x | - | - | - | - | - | - | -
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2008-02-01
| F.0 Patch |  | F.01.46 | F.01.46 | x | x | x | x | x | x | x | x | x | x | x | x | - | F.01.46 | F.01.46 | x | - | - | - | - | - | - | SB86200840
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | [7 empty] | SB86200850
2008-10-30
| F.0 Patch |  | F.01.47 | F.01.47 | x | x | x | x | x | x | x | x | x | x | x | x | - | - | - | x | - | - | - | - | - | - | FCO86200973
|  |  | (HW A/B/C) | (A.00.04/05/06)
2008-10-30
| F.0 Patch |  | - | - | x | x | x | x | x | x | x | x | x | x | x | x | - | - | - | x | - | F.01.47 | - | - | - | - | SB86201021A
(A.00.03/04/05/08/10/12)
2009-10-01
| G.0 | NLS Rev. 6.50 | G.01.07 | G.01.11 | x | x | x | x | x | x | x | x | x | x | x | x | G.01.10 | G.01.10 | G.01.10 | x | G.01.10 | G.01.10 | G.01.10 | G.01.10 | G.01.10 | G.01.10 | -
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2008-10-30
| G.0 |  | G.01.63 | G.01.66 | x | x | x | x | x | x | x | x | x | x | x | x | G.01.66 | G.01.66 | G.01.66 | x | G.01.66 | G.01.66 | G.01.67 | G.01.67 | G.01.67 | G.01.67 | SB86200885
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2009-02-10
| G.0 |  | G.01.72 | G.01.75 | x | x | x | x | x | x | x | x | x | x | x | x | G.01.72 | G.01.74 | G.01.74 | x | G.01.73 | G.01.73 | G.01.72 | G.01.72 | G.01.72 | G.01.72 | FCO86200985A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2009-04-10
| G.0 |  | - | G.01.76 | x | x | x | x | x | x | x | x | x | x | x | x | - | G.01.76 | G.01.76 | x | G.01.76 | G.01.76 | G.01.76 | G.01.76 | G.01.76 | G.01.76 | Hungrian Language only
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | FCO86201019
2009-05-10
| G.0 |  | G.01.76 | G.01.75/G.01.76 | x | x | x | x | x | x | x | x | x | x | x | x | G.01.72 | G.01.74/G.01.76 | G.01.74/G.01.76 | x | G.01.73/G.01.76 | G.01.73/G.01.76 | G.01.72/G.01.76 | G.01.72/G.01.76 | G.01.72/G.01.76 | G.01.72/G.01.76 | without/with Hungrian Language
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | FCO86201023
2009-06-22
| G.0 |  | - | - | x | x | x | x | x | x | x | x | x | x | x | x | G.01.73 | G.01.78 | G.01.78 | x | - | G.01.79 | G.01.78 | G.01.78 | G.01.78 | G.01.78 | SB86201021A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  |  | (A.00.03/04/05/08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | FCO86201056A
| 2009-10-01 | [25 empty] | SB86201024A
| G.0 |  | G.01.80 | G.01.80 | x | x | x | x | x | x | x | x | x | x | x | x | - | G.01.80 | G.01.80 | x | G.01.80 | G.01.80 |  | G.01.80 | G.01.80 | G.01.80 | SB86201123A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (A.00.04/17/20/21) | (A.00.03/04/05/08/10/12) |  | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2010-05-05
| G.0 |  | - | G.01.81 | x | x | x | x | x | x | x | x | x | x | x | x | - | G.01.81 | G.01.81 | x | - | - | - | - | - | - | SB86202791A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2018-04-10
Software H.0 is no longer supported | H.0 | NLS Rev. 6.55 | H.03.11 | H.03.11 | x | x | x | x | x | x | x | x | H.03.10 | x | x | x | H.03.05 | H.03.11 | H.03.11 | H.03.11 | H.03.09 | H.03.09 | H.03.09 | H.03.09 | H.03.09 | H.03.09 | SB86201202A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2010-09-01
| H.0 Patch |  | H.03.12 | H.03.12 | x | x | x | x | x | x | x | x | H.03.12 | x | x | x | - | H.03.12 | H.03.12 | H.03.12 | H.03.12 | H.03.13 | H.03.13 | H.03.13 | H.03.13 | H.03.13 | FCO86201194A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2010-12-01
| H.0 Patch |  | - | H.03.14 | x | x | x | x | x | x | x | x | H.03.14 | x | x | x | - | H.03.14 | H.03.14 | H.03.14 | H.03.14 | H.03.14 | H.03.14 | H.03.14 | H.03.14 | H.03.14 | -
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2011-02-02
| H.15 | NLS Rev. 6.60 | H.15.22 | H.15.27 | x | x | x | x | x | x | x | x | H.15.29 | - | - | - | H.15.22 | H.15.28 | H.15.28 | H.15.28 | H.15.27 | H.15.27 | H.15.27 | H.15.27 | H.15.27 | H.15.27 | FCO86201206A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | FCO86201230A
2011-05-05
| H.15 Patch |  | - | H.15.31 | x | x | x | x | x | x | x | x | H.15.31 | H.15.31 | H.15.31 | H.15.22 |  | H.15.31 | H.15.31 | H.15.31 | H.15.31 | H.15.31 | H.15.31 | H.15.31 | H.15.31 | H.15.31 | FCO86201250A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2011-06-06
| H.15 Patch |  | H.15.36 | H.15.36 | x | x | x | x | x | x | x | x | H.15.36 | H.15.36 | H.15.36 | H.15.34 | H.15.34 | H.15.36 | H.15.36 | H.15.36 | H.15.36 | H.15.36 | H.15.36 | H.15.36 | H.15.36 | H.15.36 | FCO86201276A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2011-09-16
| H.15 Patch |  | - | H.15.41 | x | x | x | x | x | x | x | x | H.15.41 | H.15.41 | H.15.41 | - | - | H.15.41 | H.15.41 | H.15.41 | H.15.41 | H.15.41 | H.15.41 | H.15.41 | H.15.41 | H.15.41 | FCO86201313A/FCO86201315A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2012-01-27
| H.15 Patch |  | - | H.15.45 | x | x | x | x | x | x | x | x | H.15.45 | H.15.45 | H.15.45 | - | - | H.15.45 | H.15.45 | H.15.45 | H.15.45 | H.15.45 | H.15.45 | H.15.45 | H.15.45 | H.15.45 | FCO86201371A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2012-07-27
| H.16 |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | - | - | H.16.26 | - | - | - | - | - | - | -
(A.00.09/10/11/16)
2012-01-27
| H.16 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | - | - | H.16.29 | - | - | - | - | - | - | -
(A.00.09/10/11/16)
2012-02-17
| H.15 Patch |  | - | H.15.51 | x | x | x | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86201649A, SB86201811A
(B.00.04)
2013-03-25
| H.15 Patch |  | - | H.15.53 | x | x | x | x | x | x | x | x | - | - | - | - | - | H.15.53 | H.15.53 | H.15.53 | - | - | - | - | - | - | SB86202791A
|  |  |  | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2018-04-10
| J.0 | NLS Rev. 6.65 | J.04.95 | J.04.95 | x | x | x | x | x | x | x | x | J.04.94 | J.04.94 | J.04.94 | J.04.88 | J.04.88 | J.04.95 | J.04.95 | J.04.95 | J.04.94 | J.04.94 | J.04.94 | J.04.94 | J.04.94 | J.04.94 | SB86201562A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2012-06-20
| J.0 Patch |  | J.04.96 | J.04.98 | x | x | x | x | x | x | x | x | J.04.98 | J.04.98 | J.04.98 | J.04.96 | J.04.96 | J.04.98 | J.04.98 | J.04.98 | J.04.98 | J.04.98 | J.04.98 | J.04.98 | J.04.98 | J.04.98 | FCO86201372A
|  |  | (HW A/B/C) | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2012-07-27
| J.0 Patch |  | J.04.99 | - | x | x | x | x | x | x | x | x | - | - | - | J.04.99 | J.04.99 | - | - | - | - | - | - | - | - | - | FCO86201412A
(HW A/B/C)
2012-12-04
| J.0 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | J.05.00 | J.05.00 | - | - | - | - | - | - | - | - | - | SB86201813A
2013-03-25
| J.0 Patch |  | J.05.07 | - | x | x | x | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201496A
(HW A/B/C)
2013-12-11
| J.0 Patch |  | J.05.08 | - | x | x | x | x | x | x | x | x | - | - | - | J.05.08 | J.05.08 | - | - | - | - | - | - | - | - | - | FCO86201527A
(HW A/B/C)
2014-02-26
| J.0 Patch |  | J.05.09 | - | x | x | x | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201799A
(HW A/B/C)
2017-06-13
| J.0 Patch |  | - | J.04.99 | x | x | x | x | x | x | x | x | - | - | - | - | - | J.04.99 | J.04.99 | J.04.99 | - | - | - | - | - | - | SB86202791A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2018-04-10
| J.0 Patch |  | J.05.10 | - | x | x | x | x | x | x | x | x | - | - | - | J.05.10 | J.05.10 | - | - | - | - | - | - | - | - | - | SB86202819A
(HW A/B/C)
2018-07-12
| J.1 | NLS Rev. 6.70 | - | J.10.26 | x | x | x | x | x | x | x | x | J.10.26 | J.10.26 | J.10.26 | - | - | J.10.26 | J.10.26 | J.10.26 | J.10.26 | J.10.26 | J.10.26 | J.10.26 | J.10.26 | J.10.26 | SB86201681A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2012-11-12
| J.1 Patch |  | - | J.10.30 | x | x | x | x | x | x | x | x | - | - | - | - | - | J.10.30 | J.10.30 | J.10.30 | - | - | - | - | - | - | FCO86201412A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2012-12-04
| J.1 Patch |  | - | J.10.33 | x | x | x | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86201649A, SB86201813A
(A.00.04/05/06, B.00.04)
2013-03-25
| J.1 Patch |  | - | J.10.45 | x | x | x | x | x | x | x | x | J.10.45 | J.10.45 | J.10.45 | - | - | J.10.45 | J.10.45 | J.10.45 | J.10.45 | J.10.45 | J.10.45 | J.10.45 | J.10.45 | J.10.45 | SB86201953A
|  |  |  | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2013-11-22
| J.1 Patch |  | - | J.10.49 | x | x | x | x | x | x | x | x | - | - | - | - | - | J.10.48 | J.10.48 | J.10.48 | - | - | - | - | - | - | FCO86201496A
|  |  |  | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2013-12-11
| J.1 Patch |  | - | J.10.50 | x | x | x | x | x | x | x | x | J.10.50 | J.10.50 | J.10.50 | - | - | J.10.51 | J.10.51 | J.10.51 | J.10.51 | J.10.51 | J.10.50 | J.10.50 | J.10.50 | J.10.50 | FCO86201527A
|  |  |  | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2014-02-26
| J.1 Patch |  | - | J.10.52 | x | x | x | x | x | x | x | x | J.10.52 | J.10.52 | J.10.52 | - | - | - | - | - | J.10.52 | J.10.52 | J.10.52 | J.10.52 | J.10.52 | J.10.52 | FCO86201610A
|  |  |  | (A.00.04/05/06, B.00.04) | [16 empty] | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2014-10-07
| J.1 Patch |  | - | J.10.53 | x | x | x | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201799A
(B.00.04)
2017-06-13
| J.1 Patch |  | - | J.10.53 | x | x | x | x | x | x | x | x | - | - | - | - | - | J.10.53 | J.10.53 | J.10.53 | - | - | - | - | - | - | SB86202791A
|  |  |  | (A.00.04/05/06) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2018-04-10
| J.1 Patch |  | - | J.10.54 | x | x | x | x | x | x | x | x | J.10.54 | J.10.54 | J.10.54 | - | - | J.10.54 | J.10.54 | J.10.54 | J.10.54 | J.10.54 | J.10.54 | J.10.54 | J.10.54 | J.10.54 | SB86202819A
|  |  |  | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2018-07-12
| J.2 | NLS Rev. 6.80 | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.21.03 | J.21.03 | J.21.03 | - | - | - | - | - | - | SB86201974A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2013-12-19
| J.2 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.21.19 | J.21.19 | J.21.19 | - | - | - | - | - | - | FCO86201527A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2014-02-26
| J.2 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.21.20 | J.21.20 | J.21.20 | - | - | - | - | - | - | FCO86201584B
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2014-08-12
| J.2 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.21.22 | J.21.22 | J.21.22 | - | - | - | - | - | - | FCO86201610A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2014-10-07
| J.2 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.21.23 | J.21.23 | J.21.23 | - | - | - | - | - | - | SB86202212A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2015-02-16
| J.2 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.21.24 | J.21.24 | J.21.24 | - | - | - | - | - | - | SB86202791A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2018-04-10
| J.2 Patch |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.21.25 | J.21.25 | J.21.25 | - | - | - | - | - | - | SB86202819A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2018-07-12
| J.2 (Epidemic |  | - | - | x | x | x | x | x | x | x | x | - | - | - | - | - | J.22.01 | J.22.01 | J.22.01 | - | - | - | - | - | - | IOSM86200040A
| Triage protocol) | [16 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2014-11-06
Software K.0/K.1 is no longer supported | K.0 |  | - | - | x | x | K.01.65 | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86201888A
(A.00.16)
2013-08-23
| K.0 Patch |  | - | - | x | x | K.01.70 | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201487A
(A.00.16)
2013-11-08
| K.0 Patch |  | - | - | x | x | K.01.71 | x | x | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201496A
(A.00.16)
2013-11-23
| K.1 |  | - | - | x | x | K.10.47 | x | K.10.47 | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86201961A
[7 empty] | (A.00.16) |  | (A.00.16)
2013-12-11
| K.1 Patch |  | - | - | x | x | K.10.71 | x | K.10.71 | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86201998A
[7 empty] | (A.00.16) |  | (A.00.16)
2014-02-26
| K.1 Patch |  | - | - | x | x | K.10.73 | x | K.10.73 | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201552A
[7 empty] | (A.00.16) |  | (A.00.16)
2014-03-14
| K.1 Patch |  | - | - | x | x | K.10.74 | x | K.10.74 | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201610A
[7 empty] | (A.00.16) |  | (A.00.16)
2014-10-07
| K.2 | NLS Rev. 6.85 | K.21.31 | K.21.30 | x | x | K.21.31 | x | K.21.31 | x | x | x | K.21.30 | K.21.30 | K.21.30 | K.21.31 | K.21.31 | K.21.31 | K.21.31 | - | K.21.30 | K.21.30 | K.21.30 | K.21.30 | K.21.30 | K.21.30 | SB86202101A
|  |  | (HW C) | (A.00.04/05/06, B.00.04) |  |  | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2014-08-29
| K.2 Patch |  | - | K.21.39 | x | x | K.21.39 | x | K.21.39 | x | x | x | K.21.39 | K.21.39 | K.21.39 | K.21.39 | K.21.39 | K.21.40 | K.21.40 | - | - | K.21.39 | K.21.39 | K.21.39 | K.21.39 | K.21.39 | FCO86201610A
|  |  |  | (A.00.04/05/06, B.00.04) |  |  | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  |  | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2014-10-07
| K.2 Patch |  | K.21.42 | K.21.42 | x | x | K.21.43 | x | K.21.43 | x | x | x | K.21.42 | K.21.42 | K.21.42 | K.21.42 | K.21.42 | K.21.42 | K.21.42 | - | K.21.42 | K.21.42 | K.21.42 | K.21.42 | K.21.42 | K.21.42 | SB86202212A
|  |  | (HW C) | (A.00.04/05/06, B.00.04) |  |  | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2015-02-16
| K.2 Patch |  | K.21.54 | K.21.54 | x | x | - | x | - | x | x | x | K.21.54 | K.21.54 | K.21.54 | K.21.54 | K.21.54 | K.21.54 | K.21.54 | - | K.21.54 | K.21.54 | K.21.54 | K.21.54 | K.21.54 | K.21.54 | FCO86201686B
|  |  | (HW C) | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2015-10-07
| K.2 Patch |  | K.21.57 | K.21.57 | x | x | - | x | - | x | x | x | - | - | - | K.21.56 | K.21.56 | K.21.57 | K.21.57 | - | - | - | - | - | - | - | FCO86201693A
|  |  | (HW C) | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2015-12-10
| K.2 Patch |  | - | K.21.58 | x | x | - | x | - | x | x | x | K.21.58 | K.21.58 | K.21.58 | - | - | K.21.58 | K.21.58 | - | K.21.58 | K.21.58 | K.21.58 | K.21.58 | K.21.58 | K.21.58 | FCO86201701B
|  |  |  | (A.00.04/05/06, B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2016-02-24
| K.2 Patch |  | - | - | x | x | K.21.46 | x | K.21.46 | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201725A
[7 empty] | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30)
2016-07-20
| K.2 Patch |  | K.21.60 | K.21.60 | x | x | - | x | - | x | x | x | - | - | - | - | - | K.21.60 | K.21.60 | - | - | - | - | - | - | - | FCO86201799A
|  |  | (HW C) | (B.00.04) | [13 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16)
2017-06-13
| K.2 Patch |  | - | K.21.60 | x | x | - | x | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86202791A
(A.00.04/05/06)
2018-04-10
| K.2 Patch |  | K.21.61 | K.21.61 | x | x | K.21.61 | x | K.21.61 | x | x | x | K.21.61 | K.21.61 | K.21.61 | K.21.61 | K.21.61 | K.21.61 | K.21.61 | - | K.21.61 | K.21.61 | K.21.61 | K.21.61 | K.21.61 | K.21.61 | SB86202819A
|  |  | (HW C) | (A.00.04/05/06, B.00.04) |  |  | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) |  | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2018-07-12
| L.0 | NLS Rev. 6.90 | L.00.96 | L.00.96 | x | x | L.00.97 | L.00.97 | L.00.97 | x | x | x | L.00.96 | L.00.96 | L.00.96 | L.00.96 | L.00.96 | L.00.96 | - | L.00.96 | L.00.96 | L.00.96 | L.00.96 | L.00.96 | L.00.96 | L.00.96 | SB86202268A
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2015-07-10
| L.0 Patch |  | L.01.05 | L.01.05 | x | x | - | - | - | x | x | x | - | - | - | L.00.97 | L.00.97 | L.01.05 | - | L.01.05 | - | - | - | - | - | - | FCO86201693A
|  |  | (HW C) | (B.00.04) | [13 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2015-12-10
| L.0 Patch |  | - | L.01.06 | x | x | L.01.08 | L.01.08 | L.01.08 | x | x | x | L.01.08 | L.01.08 | L.01.08 | L.01.06 | L.01.06 | L.01.06 | - | L.01.06 | L.01.06 | L.01.08 | L.01.08 | L.01.08 | L.01.08 | L.01.08 | FCO86201701B
|  |  |  | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2016-02-24
| L.0 Patch |  | L.01.09 | L.01.09 | x | x | - | - | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86202447A
|  |  | (HW C) | (B.00.04)
2016-04-20
| L.0 Patch |  | - | - | x | x | L.01.10 | L.01.10 | L.01.10 | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201725A
[7 empty] | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30)
2016-07-20
| L.0 Patch |  | - | - | x | x | L.01.12 | L.01.12 | L.01.12 | x | x | x | L.01.09 | L.01.09 | L.01.09 | - | - | - | - | - | - | - | - | - | - | - | SB86202484A
[7 empty] | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30)
2016-09-14
| L.0 Patch |  | L.01.15 | L.01.15 | x | x | - | - | - | x | x | x | - | - | - | - | - | L.01.15 | - | L.01.15 | - | - | - | - | - | - | FCO86201799A
|  |  | (HW C) | (B.00.04) | [13 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2017-04-20
| L.0 Patch |  | L.01.19 | L.01.19 | x | x | L.01.18 | L.01.18 | L.01.18 | x | x | x | L.01.17 | L.01.17 | L.01.17 | L.01.17 | L.01.17 | L.01.19 | - | L.01.19 | L.01.17 | L.01.17 | L.01.17 | L.01.17 | L.01.17 | L.01.17 | FCO86201794A
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2017-07-28
| L.0 Patch |  | L.01.20 | L.01.21 | x | x | L.01.21 | L.01.21 | L.01.21 | x | x | x | L.01.21 | L.01.21 | L.01.21 | L.01.20 | L.01.20 | L.01.21 | - | L.01.21 | L.01.21 | L.01.21 | L.01.21 | L.01.21 | L.01.21 | L.01.21 | SB86202819A, SB86202836A
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2018-07-12
| L.0 Patch 12 |  | L.01.22 | L.01.22 | x | x | L.01.23 | L.01.23 | L.01.23 | x | x | x | L.01.22 | L.01.22 | L.01.22 | L.01.22 | L.01.22 | L.01.22 | - | L.01.22 | L.01.22 | L.01.22 | L.01.22 | L.01.22 | L.01.22 | L.01.22 | FCO86201902A
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2020-02-03
| L.1 | NLS Rev. 6.100 | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | L.10.75 | - | L.10.75 | - | - | - | - | - | - | SB86202364A
[18 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2016-02-24
| L.1 Patch |  | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | L.10.77 | - | L.10.77 | - | - | - | - | - | - | SB86202447A
[18 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2016-04-20
| L.12 |  | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | L.12.46 | - | L.12.46 | - | - | - | - | - | - | SB86202476B
[18 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2016-09-14
| L.12 |  | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86202507A
2016-11-08
| L.12 Patch |  | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | L.12.49 | - | L.12.49 | - | - | - | - | - | - | FCO86201799A
[18 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2017-04-20
| L.12 Patch |  | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | L.12.51 | - | L.12.51 | - | - | - | - | - | - | FCO86201794A
[18 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2017-07-28
| L.12 Patch |  | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | L.12.52 | - | L.12.52 | - | - | - | - | - | - | SB86202665A
[18 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2017-09-26
| L.12 Patch |  | - | - | x | x | - | - | - | x | x | x | - | - | - | - | - | L.12.57 | - | L.12.57 | - | - | - | - | - | - | SB86202819A
[18 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2018-07-12
| M.0 | NLS Rev. 6.110 | M.03.00 (Build 347) | M.03.00 (Build 347) | x | M.03.00 (Build 346) | M.03.00 (Build 346) | M.03.00 (Build 346) | M.03.00 (Build 346) | x | x | x | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | M.03.00 (Build 347) | SB86202599A
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2017-06-13
| M.0 Patch 1 |  | M.03.01 (Build 86) | M.03.01 (Build 86) | x | - | M.03.01 (Build 85) | M.03.01 (Build 85) | M.03.01 (Build 85) | x | x | x | - | - | - | - | - | M.03.01 (Build 86) | M.03.01 (Build 86) | M.03.01 (Build 86) | M.03.01 (Build 85) | M.03.01 (Build 85) | M.03.01 (Build 85) | M.03.01 (Build 85) | M.03.01 (Build 85) | M.03.01 (Build 85) | FCO86201794A
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2017-07-28
| M.0 Patch 2 |  | - | M.03.02 (Build 22) | x | M.03.02 (Build 27) | - | - | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201806A
|  |  |  | (B.00.04) | [22 empty] | SB86202682B
2017-09-26
| M.0 Patch 3 |  | - | - | x | - | - | - | - | x | x | x | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | - | - | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | M.03.03 (Build 3) | FCO86201809A
[18 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2017-10-23
| M.04 | NLS Rev. 6.125 | M.04.00 (Build 147) | M.04.00 (Build 149) | M.04.00 (Build 147) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | x | x | x | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 147) | M.04.00 (Build 147) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | M.04.00 (Build 149) | SB86202815A, SB86202819A,
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | SB86202836A, FCO86201840A
2018-07-12
| M.04 Patch 1 |  | - | - | - | - | - | - | - | x | x | x | M.04.01 (Build 8) | M.04.01 (Build 8) | M.04.01 (Build 8) | - | - | - | - | - | - | - | - | - | - | - | -
2019-01-15
| M.04 Patch 3 |  | - | - | M.04.03 (Build 5) | - | - | - | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201872A
2019-03-27
| M.04 Patch 4 |  | - | M.04.04 (Build 7) | - | - | - | - | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201872A
(B.00.04)
2019-04-04
| M.04 Patch 5 |  | M.04.05 (Build 17) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | x | x | x | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | M.04.05 (Build 19) | FCO86201884A
|  |  | (HW C) | (B.00.04) |  |  | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (A.00.09/10/11/16) | (>= A.00.20) | (A.00.08/10/12) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23) | (B.00.04/05/06/18/22/23)
2019-09-13
| M.04 Patch 6 |  | M.04.06 (Build 3) | - | - | - | - | - | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | FCO86201884A
(HW C)
2019-09-13
| M.04 Patch 7 |  | - | - | M.04.07 (Build 6) | M.04.07 (Build 6) | - | - | - | x | x | x | - | - | - | - | - | - | - | - | - | - | - | - | - | - | SB86203035A
2020-02-03
| N.0 | NLS Rev. 6.135 | x | x | N.00.00 (Build 1452) | N.00.00 (Build 1452) | N.00.00 (Build 1452) | x | N.00.00 (Build 1452) | N.01.00 (Build 181) | N.01.00 (Build 181) | N.01.00 (Build 181) | N.00.00 (Build 1452) | N.00.00 (Build 1452) | x | N.00.00 (Build 1452) | N.00.00 (Build 1452) | N.00.00 (Build 1452) | x | N.00.00 (Build 1452) | x | x | x | x | x | x | SB86202991A, SB86202995A
[7 empty] | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2019-10-11
| N.0 Patch 1 |  | x | x | N.00.01 (Build 13) | N.00.01 (Build 13) | N.00.01 (Build 13) | x | N.00.01 (Build 13) | N.01.01 (Build 13) | N.01.01 (Build 13) | N.01.01 (Build 13) | N.00.01 (Build 13) | N.00.01 (Build 13) | x | N.00.01 (Build 13) | N.00.01 (Build 13) | N.00.01 (Build 13) | x | N.00.01 (Build 13) | x | x | x | x | x | x | FCO86201902A
[7 empty] | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2020-02-03
| N.0 Patch 2 |  | x | x | - | - | - | x | - | - | - | N.01.02 (Build 2) | - | - | x | - | - | - | x | - | x | x | x | x | x | x | FCO86201908A
2020-06-10
| N.0 Patch 3 | NLS Rev. 6.140 | x | x | N.00.03 (Build 79) | N.00.03 (Build 79) | - | x | - | N.01.03 (Build 93) | N.01.03 (Build 93) | N.01.03 (Build 93) | - | - | x | - | - | - | x | - | x | x | x | x | x | x | For limited distribution only
2020-10-14
| N.0 Patch 4 |  | x | x | N.00.04 (Build 108) | N.00.04 (Build 108) | N.00.04 (Build 108) | x | N.00.04 (Build 108) | N.01.04 (Build 87) | N.01.04 (Build 87) | N.01.04 (Build 87) | N.00.04 (Build 108) | N.00.04 (Build 108) | x | N.00.04 (Build 108) | N.00.04 (Build 108) | N.00.04 (Build 108) | x | N.00.04 (Build 108) | x | x | x | x | x | x | SB86203146A
[7 empty] | (A.00.16/20/22/24/30) |  | (A.00.16/20/22/24/30) | [8 empty] | (A.00.09/10/11/16) |  | (A.00.09/10/11/16)
2021-01-22
This overview shows all upgradeable SW revisions released with the Intellivue family. In the Support Tool some of them are no longer available because they were replaced by newer revisions. Please use always the latest Support Tool to service the devices.
Note on E8: The M3001A has different HW versions which can be distinguished by the Serial Number Prefix and the option string. HW A: #A01 DE227, DE441 HW B: #A01 DE512; #A02 DE441, DE512 HW C: DE345, DE420, DE468, DE522, DE563, DE652 #A01 >=DE610; #A02 >=DE610; #A03 >=DE632
Note on U8: Product discontinued and no longer sold Planned End of Support: 31-Dec-2022
Note on W8: Product discontinued and no longer sold Planned End of Support: 31-Dec-2022
Note on X8: Product discontinued and no longer sold Planned End of Support: 31-Dec-2022
Note on Y8: Product discontinued and no longer sold End of Support: 31-Dec-2020
Note on Z8: Product discontinued and no longer sold End of Support: 31-Dec-2020
Note on AA8: Product discontinued and no longer sold End of Support: 31-Dec-2020
Note on AB8: Product discontinued and no longer sold End of Support: 31-Dec-2020
Note on E49: No longer included in Support Tool, see FCO86200673
Note on E52: No longer included in Support Tool, see FCO86200673
Note on E56: No longer included in Support Tool, see FCO86200673
Note on E68: No longer included in Support Tool, see FCO86200673
Note on E81: see FCO86200673
Note on V174: MP5SC only H.16.xx enables IntelliVue Guardian SW compatibility
Note on V177: MP5SC only H.16.xx enables IntelliVue Guardian SW compatibility
Note on C212: B300-2018-06-23854
Note on C242: B300-2018-06-23854
Note on T264: Software version J.22.01 does not support connectivity to PIIC iX or IntelliVue Guardian Software (IGS). If software version J.22.01 is loaded on to an MP5 with option C15, this option will be deactivated to prevent connection to PIIC iX. MP5 and MP5SC with J.22.01 are not allowed to connect to IGS. Monitors with software version J.22.01 can only be used as standalone triage devices. Unless you have a need for demonstrating or installing a monitor with the Epidemic Triage protocol, do not upgrade any monitor to J.22.01.
Note on U264: Software version J.22.01 does not support connectivity to PIIC iX or IntelliVue Guardian Software (IGS). If software version J.22.01 is loaded on to an MP5 with option C15, this option will be deactivated to prevent connection to PIIC iX. MP5 and MP5SC with J.22.01 are not allowed to connect to IGS. Monitors with software version J.22.01 can only be used as standalone triage devices. Unless you have a need for demonstrating or installing a monitor with the Epidemic Triage protocol, do not upgrade any monitor to J.22.01.
Note on V264: Software version J.22.01 does not support connectivity to PIIC iX or IntelliVue Guardian Software (IGS). If software version J.22.01 is loaded on to an MP5 with option C15, this option will be deactivated to prevent connection to PIIC iX. MP5 and MP5SC with J.22.01 are not allowed to connect to IGS. Monitors with software version J.22.01 can only be used as standalone triage devices. Unless you have a need for demonstrating or installing a monitor with the Epidemic Triage protocol, do not upgrade any monitor to J.22.01.
Note on C317: B300-2018-06-23854
Note on C344: B300-2018-06-23854
Note on C347: B300-2019-12-25960
Note on C371: B300-2018-06-23854
Note on C386: B300-2018-06-23854
Note on C389: B300-2018-11-24364
Note on C392: B300-2019-03-24885
Note on C395: B300-2019-03-24910
Note on E396: For limited distribution only
Note on C398: B300-2019-05-25138 (M3001A/AL only) B300-2019-08-25441
Note on C401: B300-2019-08-25441
Note on C404: B300-2019-12-25960
Note on C407: B300-2019-06-25172 (N.0) B300-2016-07-21350 (N.01)
Note on C410: B300-2019-12-25960
Note on C413: B300-2020-03-26258
Note on C416: B300-2020-08-26845
Note on C419: B300-2020-12-27271

## FW IntelliVue
Intellivue FW Revisions History
NOTE: On the Multi measurement Module HW A the FW is fix and cannot be upgraded
The M3001A has different HW versions which can be distinguished by the Serial Number Prefix and the option string.
HW A: #A01 DE227, DE441
HW B: #A01 DE512 ; #A02 DE441, DE512
HW C: #A01 >= DE610; #A02 >=DE610; #A03>=DE632; #A04 DE907
On HW B and HW C, the SpO2 and the ECG FW are upgradeable. Please see following tables for SpO2 and ECG FW.
To perform the FW upgrade the MMS must have SW Revision C.0 or higher and connected to an Intellivue Monitor. On an M3046A the FW upgrade is not supported.
ECG
FW Rev | Related Products |  |  | Comment
C.00.13 |  |  |  | The HW and its interface do not allow a FW upgrade.
C.01.19
D.01.70 | M3001A |  |  | Initial revision for HW B
D.01.76 | M3001A |  |  | FCO86200612 (In rare cases, physiological spikes or notches in the ECG might trigger the ECG pacer detection algorithm. This can lead to a distorted ECG wave) Retrofit on failure upgrade to D.01.76
D.01.77 | M3001A |  |  | FCO86200624 (MMS with ECG firmware revision D.01.76 might loose the ECG signal when the power to the MMS is interrupted for a short period of time (about 3 seconds). In this case, a flat ECG wave would be shown and no ECG-related numeric would be available (HR, ST,…). In addition, an INOP “ECG Noisy Signal” will be displayed.) Retrofit on failure upgrade to D.01.77
D.01.78 | M3001A/M8105A |  |  | FCO86200654: MMS with ECG firmware revision D.01.77 or lower might not correctly detect all pace pulses when less than 5 ECG electrodes are attached to the patient. In rare cases, this could lead to an incorrect heart rate displayed on the monitor or missing heart rate alarms. Retrofit on failure upgrade to D.01.78
D.01.89 | M3001A, M8105A |  |  | FCO86200733: Noise and Pacer improvements with new ECG Firmware D.01.89 Retrofit on failure upgrade to D.01.89
D.02.02 | M3001A/M3002A, M8102A/M8105A |  |  | 6ld Support ECG HW revision D.01.01 required NOTE: ONLY RUNNING WITH SW F.0 and higher
D.02.05 | M3001A/M3002A, M8102A/M8105A |  |  | The ECG Firmware revision D.02.05 provides an improved pace pulse detection specificity especially in situations with high common mode noise. NOTE: ONLY RUNNING WITH SW F.0 and higher. FCO86201126A
E.01.10 | M3001A (starting with SN Prefix DE468) |  |  | Initial revision for ECG HW E.01.01
E.01.16 | M3001A (starting with SN Prefix DE522), M3002A (starting with SN Prefix DE535), M8102A (starting with SN Prefix DE536), M8105A (starting with SN Prefix DE537) |  |  | The ECG Firmware revision E.01.16 provides an improved handling of unusual high electrode to skin impedances (> 200 kOhm) FCO86201628A
E.01.22 | M3001A/M3002A, M8102A/M8105A |  |  | SB86202447A
F.01.30 | 867030 (X3), 867033 (MX100), 867036 (MMX) |  |  | Initial revision
SpO2
Option |  | SpO2 DSP FW Rev |  | Related Products |  | Comment
#A01 FAST SpO2 |  | A.01.04 |  | M3001A |  | The HW and its interface do not allow a FW upgrade.
|  | A.01.35/A.01.37 |  | M1020B
#A01/SP1 FAST SpO2 |  | (A.01.41) -> A.01.46 |  | M3001A, M1020B |  | FCO86200673: MMS occasionally displays unexpected pulse oximetry readings Mandatory upgrade to A.01.46
(A.01.42) -> A.01.46
|  | A.01.46 |  | M1020B, M3001A/M3002A, M8102A/M8105A/5AT |  | ---
|  | A.01.47 |  |  |  | ---
|  | A.01.48 |  |  |  | ---
|  | A.04.12 for HW B.03.00 |  | M3001A/2A M8102A/5A/5AS/5AT |  | ---
|  | A.04.13 for HW B.03.00 |  |  |  | SB86201501A: Improved identification of incorrectly applied sensors to better support sensors which are working at their specification limits
|  | A.04.16 for HW B.03.00 |  | M3001A/2A M8102A/5A/5AS/5AT, M1020B |  | SB86201696A: M3001/2A, M8102/5A/AT/AS Philips FAST SPO2 New Firmware available
|  | A.04.17 |  | M3001A/2A M8102A/5A/5AS/5AT |  | ---
#A02 Nellcor Oximax Compatible |  | (A.01.41) -> A.01.46 |  | M3001A, M1020B |  | FCO86200673: MMS occasionally displays unexpected pulse oximetry readings Mandatory upgrade to A.01.46
(A.01.42) -> A.01.46
|  | A.01.46 |  | M1020B, M3001A |  | ---
|  | A.01.47 |  |  |  | ---
#A03/SP3 MASIMO |  | Masimo SET FW |  | M3001A/M3002A |  | Not upgradeable by Philips
#A04/SP4 Nellcor Technology |  | (A.18.49) -> A.18.65 |  | M3001A/2A, M8102A/5A/5AS/5AT |  | FCO86201140A: SpO2 Nellcor Oximax™ shows intermittend SpO2 sensor malfunction INOP Proactive upgrade to A.18.65
|  | A.18.64 |  |  |  | Contains Fix for Sensor detection (no further upgrade available)
|  | A.18.65 |  |  |  | HW Rev. N.01.00; Contains Fix for Sensor detection
|  | A.18.81 |  |  |  | HW Rev. N.01.00 To upgrade to FW revision A.18.81 FW ≥ A.18.65 is required; FW A.18.64 does not need an upgrade
#A05/SP5 Masimo rainbow Set® |  | Masimo rainbow Set® FW |  | M3001AL/M3002A, M8102A/5A/5AS, 867030, 867033 |  | Not upgradeable by Philips
Inv. Pressure/Temperature
FW Rev | Related Products |  |  | Comment
D.00.27 | 867030 (X3), 867033 (MX100), 867036 (MMX), 867039/40/41 Extensions |  |  | Initial revision
D.00.28 | 867030 (X3), 867033 (MX100), 867036 (MMX), 867039/40/41 Extensions |  |  | ---
NBP
FW Rev | Related Products |  |  | Comment
A.00.44 | 867030 (X3), 867033 (MX100) |  |  | Initial revision
A.00.45 | 867030 (X3), 867033 (MX100) |  |  | FCO86201806A (It was observed that the X3 will intermittently only show NBP mean values.) Proactive upgrade to A.00.45
A.00.50 | 867030 (X3), 867033 (MX100), 867036 (MMX) |  |  | FCO86201840A: IVPM - New Patch for IntelliVue Software M.0 Retrofit on Failure upgrade to A.00.50
A.00.51 | 867030 (X3), 867033 (MX100), 867036 (MMX) |  |  | SB86202989A: NBP enhancements that will apply if artifact suppression "high" is selected
CO2
Option | SN Prefix | FW Rev | Related Products |  | Comment
Basic option (w/o P/T) #C06 (w P/T) | DE020xxxxx (HW A) | O.01.03 |  |  | Not upgradeable
|  | O.01.10 | M3015A |  | Internal HW changes required new FW
|  | O.01.30 | M3015A |  | Additional features in FW: Suppresses automatic Zeroing for 5 min., Switching Pump off for Lavage procedure
| DE435xxxxx (HW B) | P.01.23 | M3015A |  | New internal HW structure Includes modification in the internal communication
P.01.24
|  | P.01.26 |  |  | FCO86200859: IVPM NEW PATCH IntelliVue SW Rev F.0 Release and new CO2 FW Service Recommended upgrade
|  | P.01.30 |  |  | SB86200914A: M3015A, New Firmware (P.01.30) for M3015A with Rev B HW
|  | P.01.31 | M3015A, M8105A/5AS |  | SB86201129A: M8105A M3015A New FW P.01.31 available
|  | P.01.32 |  |  | SB86201587A: M8105A, M3015A/B New FW rev. P.01.32/P.01.46 available
| DE138xxxxx (HW C) | P.01.46 | M3015A/B |  | HW Rev. Q.01, SB86201587A: M8105A, M3015A/B New FW rev. P.01.32/P.01.46 available
SO2
FW Rev | Related Products |  | Comment
A.00.11 | M1011A |  | Initial revision
A.00.12 |  |  | SB86201128A: M1011A New FW Revision A.00.12 available
Spirometry
FW Rev | Related Products |  | Comment
A.01.13 | M1014A |  | Initial revision
EEG
FW Rev | Related Products |  | Comment
A.06.09 | M1027B |  | SB86202275A: M1027B - aEEG and BSR Functionality is now available
A.07.00 |  |  | SB86203035A: Improved “EEG High Impedance” INOP handling for electrodes that are actually not being used
NMT
FW Rev | Related Products |  | Comment
A.00.10 | 865383 |  | SB86201879A 865383 NMT Module new firmware FW A.00.10 available
802.11 WLAN for M80xxA, M8105A/5AS
FW Rev | Related Products |  | Comment
A.01.12 | (M80xxA #J35) |  | FCO86200744: Roaming performance improvement of IntelliVue 802.11 wireless LAN adapter Retrofit on failure upgrade to A.01.16
A.01.16 | (M80xxA, M8105A #J35) |  | FCO86200817: Duplicate IP address error causing reboot of IntelliVue Information Center. Retrofit on failure upgrade to A.01.18
A.01.18 |  |  | ---
A.02.23 |  |  | Replaced by A.02.24
A.02.24 |  |  | Replaced by A.02.25
A.02.25 |  |  | Replaced by A.03.05
A.03.05 | (M80xxA, M8105A/5AS #J35) |  | Enterprise authentication available SB86201267A: IVPM New WLAN firmware available
A.03.06 |  |  | FCO86201316A: IVPM Multicast issue and failed WLAN Enterprise Authentication with WLAN FW A.03.05
A.03.08 |  |  | SB86201781A: IVPM: new WLAN FW A.03.08 available
A.03.09 |  |  | ---
A.03.11 |  |  | SB86203067A: IVPM 802.11 (WLAN) Security Vulnerability Patch Firmware A.03.11 available
802.11 WLAN for M3002A/M8102A, M8105A/5AS, MX600-800
Device FW Rev | Related Products |  | Comment
A.00.19 | M3002A/M8102A |  | Initial revision
A.00.20 | M3002A/M8102A, MX800 |  | Improved roaming.
A.01.09 | M3002A/M8102A, MX600-800 |  | Enterprise authentication available SB86201267A: IVPM New WLAN firmware available
B.00.17 | M3002A/M8102A/5A/5AS, MX600-800 |  | HW Rev. B.00.01
B.00.24 |  |  | FCO86201701B
B.00.26 |  |  | ---
B.00.27 |  |  | ---
B.00.28 |  |  | FCO86201794A
B.00.31 |  |  | SB86202836A
IIT
IIT Telemetry board consists of two parts: IIT System board with the Device FW and the Radio Modul with the Radio FW.
The following tables show the valid combination of both.
IIT 1.4 GHz (US only)
Device FW Rev | Radio FW Rev | Related Products |  | Comment
A.00.10 | A.00.44 | M8001A/2A, M8105A #J45, M2638A |  | Initial revision
A.00.10 | A.00.52 |  |  | HW revision change, requires new Radio FW
A.00.17 | A.00.52 | M8001A/2A, M8102A/5A #J45, M2638A |  | Location Tracking added
A.00.17 | A.00.57 |  |  | Roaming improvement
A.00.19 | A.00.57 | M8001A/2A, M8102A/5A, MX600-800 #J45, M2638A |  | ---
A.00.20 | A.00.57 | M8001A/2A, M8102A/5A, MX600-800 #J45, M2638A |  | For Radio HW Rev. A.00.0x
| A.01.73 |  |  | For Radio HW Rev. A.01.01
IIT 2.4 GHz (Worldwide)
Device FW Rev | Radio FW Rev | Related Products |  | Comment
A.00.17 | A.01.04 | M8001A/2A, M8102A/5A #J47, M2640A |  | Initial revision
A.00.17 | A.01.06 |  |  | Roaming improvement
A.00.19 | A.01.06 | M8001A/2A, M8102A/5A #J47, M2640A |  | (for Radio HW Rev. A.00.0x/C.00.01)
| A.01.51 | M8001A/2A, M8102A/5A, MX600-800 #J47, M2640A |  | Radio board hardware change requires new FW (for Radio HW Rev. D.00.01)
A.00.20 | A.01.06 | M8001A/2A, M8102A/5A #J47, M2640A |  | For Radio HW Rev. A.00.0x/C.00.01
| A.01.51 | M8001A/2A, M8102A/5A, MX600-800 #J47, M2640A |  | Radio board hardware change requires new FW (for Radio HW Rev. D.00.01)
SRR
FW Rev | Related Products |  | Comment
0.12 | M3002A, M8102A/5A/5AS/5AT |  | Initial revision
A.00.83 |  |  | SB86201191A: M3002A M8102A M8105A AS T New SRR FW Revision A.00.83
A.10.80 | M3002A, M8102A/5A/5AS/5AT, MX600 - 800 |  | SB86201567A: M3002A/M8102A/M8105A/AS/AT/MX6/7/800 – New SRR FW Revision A.10.80 available
B.04.00 |  |  | ---
B.02.60 | 866060/62/64/66/74 |  | Initial revision
D.00.42 |  |  | SB86202759A: M2702/3/4/5A New Avalon CL Wide Range/CL Fetal & Maternal pod/Avalon SW Rev. L.3 available

## Compatibility IntelliVue
Intellivue Monitor Compatibility
O | compatible |  | X | not compatible
Monitor Software1) | MMX (867036)
| M.04 | N.0
< J.x | X | X
J.x | O | O
K.2 | O | O
L.0 | O | O
M.0x | O | O
N.0 | O | O
Note: It is recommended the MMX and the host monitor are on the same major software version, otherwise release-specific functionality might not be available
1) MMX #SP5 (Masimo rainbow SET® SpO2) requires a host monitor with software L.0 or higher
Monitor Software1) | X3 (867030)
| M.0x | N.0
< J.x | X | X
J.x | O2) | O2, 3)
K.2 | O2) | O2, 3)
L.0 | O2) | O2, 3)
M.0x | O | O3)
N.0 | O | O4)
Note: When used in companion mode, it is recommended the X3 and the host monitor are on the same major software version, otherwise release-specific functionality might not be available
1) X3 #SP5 (Masimo rainbow SET® SpO2) requires a host monitor with software L.0 or higher
2) Requires X3 with software M.03.02 (or higher) and a host monitor with a specific patch level, see SB86202682B
3) Once one of the Inv. Pressure or Temp labels that were introduced with software N.0 as part of the Full Label Set is used, the X3 requires a host monitor with software N.0 or higher
4) The Inv. Pressure or Temp labels that were introduced with software N.0 as part of the Full Label Set are not supported with MX700/800
Hemo App Software | X3 (867030)
| M.0x | N.0
< 1.2.2 | X | X
>= 1.2.2 | X | O
Notes: X3 requires option C99 to to connect to the hemo system. For the continous workflow, option T9X is required in addition to C99. The X3 also requires the configuration of a specific hemo profile.
Monitor Software | MMS (M3001A/M3001AL)1)
HW A #A01 (SN Prefix (SNP) DE227, DE441)
HW B (#A01 SNP DE512) (#A02 SNP DE441, DE512)
HW C (#A01, #A02, #A03, #A042) SNP DE610, DE632, DE717, DE907, DE345, DE420, DE468, DE522, DE563, DE652)
| B.0 | B.1 | C.0 | D.0 | E.0 | F.0 | G.0 | H.15 | J.0 | K.2 | L.0 | M.0x
A.2 | O | O | O | O | O | O | O | X | X | X | X | X
B.0 | O | O | O | O | O | O | O | O | O5) | O5) | O5) | O5)
B.1 | O | O | O | O | O | O | O | O | O | O | O | O
C.0 | O | O | O | O | O | O | O | O | O | O | O | O
D.0 | O | O | O | O | O | O | O | O | O | O | O | O
E.0 | X | X | O | O | O | O | O | O | O | O | O | O
F.0 | X | X | O | O | O | O | O | O | O | O | O | O
G.0 | X | X | X | O | O | O | O | O | O | O | O | O
H.15 | X | X | X | X | X | O | O | O | O | O | O | O
J.x | X | X | X | X | X | O | O | O | O | O | O | O
K.2 | X | X | X | X | X | X | O3) | O | O | O | O | O
L.x | X | X | X | X | X | X | O3) | O | O | O | O | O
M.0x | X | X | X | X | X | X | O3) | O | O | O | O | O
N.0 | X | X | X | X | X | X | O3) | O4) | O | O | O | O
Note: It is recommended the MMS and the host monitor are on the same major software version, otherwise release-specific functionality might not be available
1) M3001AL (with Masimo rainbow SET® SpO2) is based on HW C; requires the MMS and a host monitor with software L.0 or higher
2) M3001A #A04 must not be downgraded below F.0
3) Not with MX400/430/450/500/550/750/850
4) Not with MX750/850
5) NBP repeat times of 1 h and above are shown as question marks only
Monitor Software1) | X2 (M3002A)
HW A (SN Prefix DE736, DE836, DE950, DE037)
HW B (SN Prefix DE315, DE422, DE466, DE506, DE535, DE575, DE654)
| F.0 | G.0 | H.15 | J.x | K.2 | L.0 | M.0x
F.0 | O2) | O2, 3) | O2, 3) | X | X | X | X
G.0 | O2) | O2) | O2) | O2) | O2) | O2) | O2)
H.15 | O2) | O2) | O | O | O | O | O
J.x | X | O2) | O | O | O | O | O
K.2 | X | O4) | O | O | O | O | O
L.0 | X | O4) | O | O | O | O | O
M.0x | X | O4) | O | O | O | O | O
N.0 | X | O4) | O5) | O | O | O | O
Note: When used in companion mode, it is recommended the X2 and the host monitor are on the same major software version, otherwise release-specific functionality might not be available
1) X2 #A05 (Masimo rainbow SET® SpO2) requires the X2 and a host monitor with software L.0 or higher
2) CL Sensors not supported
3) Not compatible if TAAP is used
4) Not with MX400/430/450/500/550/750/850
5) Not with MX750/850
MP2 (M8102A)1)
HW A (SN Prefix (SNP) DE736, DE836, DE950, DE037)
HW B (SN Prefix (SNP) DE315, DE423, DE467, DE507, DE536, DE576, DE655)
F.0 | G.0 | H.15 | J.x | K.2 | L.0 | M.0x
1) Except Military Bundle, see SB86201649A
Monitor Software | FMX-4 (866468)1)
N.0
< N.0 | X
N.0 | O
1) Supported with MX750/850 only
Monitor Software | FMS-8 (M8048A), FMS-4 (865243, starting with H.15)1)
| B.0 | B.1 | C.0 | D.0 | E.0 | F.0 | G.0 | H.15 | J.0 | K.2 | L.0 | M.0x | N.0
B.02) | O | O | O | O | O | X | X | X | X | X | X | X | X
B.12) | O | O | O | O | O | X | X | X | X | X | X | X | X
C.0 | O | O | O | O | O | X | X | X | X | X | X | X | X
D.03) | O | O | O | O | O | X | X | X | X | X | X | X | X
E.0 | X | X | O | O | O | O | X | X | X | X | X | X | X
F.0 | X | X | O | O | O | O | X | X | X | X | X | X | X
G.0 | X | X | X | X | X | X | O | X | X | X | X | X | X
H.15 | X | X | X | X | X | X | X | O | X | X | X | X | X
J.x | X | X | X | X | X | X | X | X | O | X | X | X | X
K.2 | X | X | X | X | X | X | X | X | X | O | X | X | X
L.0 | X | X | X | X | X | X | X | X | X | X | O | X | X
M.0x | X | X | X | X | X | X | X | X | X | X | X | O | X
N.0 | X | X | X | X | X | X | X | X | X | X | X | X | O4)
1) Supported with MP60 - 90 and MX600 - 800 only
2) MP60/70 requires C40 to support FMS
3) It is recommended that any FMS connected to an Rel D.0 or higher MP90 system should have at least revision C.0 or higher
4) Supported with MX700/800 only
| 1x FMS-4 | 1x FMS-8 | 2x FMS-4 | 2x FMS-8 | 1x FMS-4 + 1x FMS-8 | 1x FMX-4 | 2x FMX-4
MP5-MP50 | X | X | X | X | X | X | X
MP60/70/80 | O | O | X | X | X | X | X
MP90 | O | O | O | O | O | X | X
MX400/430/450 | X | X | X | X | X | X | X
MX500/550 | X | X | X | X | X | X | X
MX600/700 | O | O1, 2) | X | X | X | X | X
MX800 | O | O | O | O | O | X | X
MX750 | X | X | X | X | X | O | O
MX850 | X | X | X | X | X | O | O
1) Four (4) modules only, plus Recorder
2) With software >= M.04.00 and option H30 (commercial option H32), 8 modules in total and up to 3 IntelliBridge EC10 modules are supported
Monitor Software | MP5 (M8105A) when used as a companion
| F.0 | G.0 | H.15 | J.x | >= J.2
F.0 | O1) | O1, 2) | O1, 2) | X | X
G.0 | O1) | O1) | O1) | O1) | X
H.15 | O1) | O1) | O | O | X
J.x | X | O1) | O | O3) | X
>= K.2 | X | X | X | X | X
1) CL Sensors not supported
2) Not compatible if TAAP is used
3) MP5 companion mode is no longer supported with software revision J.2 and higher
Monitor Software | M3012A, M3014A, M3015A1)
all | O
1) X3 and MMX not with M3015A HW B, see SB86202966A
Monitor Software | M3015B in combination with
| MMS with SW >= H.0 | X2 with SW >= H.0
B.0 | O | X
B.1 | O | X
C.0 | O | X
D.0 | O | X
E.0 | O | X
>= F.0 | O | O
Measurement Server Software | 8670391) Hemodynamic Extension | 8670401) Capnography Extension | 8670411) Microstream Extension
< A.2 | X | X | X
>= A.2 - B.x | O | X | X
>= C.0 | O | O2) | X
>= D.0 - G.0 | O | O | X
>= H.0 | O | O | O
1) Dual Pressure via Y-cable only available with MMX, X3/MX100 and host monitor >= M.0x
2) Mainstream CO2 only
Measurement Description | Module Number | MP20/MP30 (No internal slots, no FMS support) |  | MP40/MP50 (No FMS support) |  | MP60/MP70 |  | MX400/MX430/MX450 (No internal slots, no FMS support) |  | MX500/MX550 (No FMS support) |  | MX750 (Two FMX-4 are supported) |  | MX850 (Two FMX-4 are supported) |  | MX600/MX700 |  | MP80/MP90/MX800 (With MP90/MX800 two FMS are supported) |  | Comments
|  | Maximum number of supported modules/measurements in total |  |  |  |  | Maximum number of supported measurements from FMS/FMXn)
Maximum number of supported measurements from Multi Measurement Module/Measurement Extension
Maximum number of supported measurements from Internal Slots
Pressure Press with analog out | M1006B M1006B #C01 | 2/3a) | x | 3/4b) | x | 6 | 5 | 2/3c) | x | 2/4d) | x | 8 | 8 | 8 | 8 | 2/4d) | 4 | 6e) | 6e)
|  |  | 3 |  | 4 |  | 4 |  | 4 |  | 4 |  | 4 |  | 4 |  | 4 |  | 4
|  |  | x |  | 2 |  | 1 |  | x |  | 3 |  | x |  | x |  | x |  | x
C.O. C.C.O. (PiCCO) | M1012A M1012A #C10 | 1 | x | 1 | x | 1 | 1 | 1f) | x | 1f) | x | 1 | x | 1 | x | 1 | 1 | 1 | 1 | For SW ver. >= E.0 PiCCO plus available
|  |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1
|  |  | x |  | 1 |  | 1 |  | x |  | 1 |  | x |  | x |  | x |  | x
Spirometry | M1014A | x | x | 1 | x | 1 | 1 | x | x | 1 | x | x | x | x | x | 1 | 1 | 1 | 1 | SW ver. >= D.0 required M1014A is not supported with FMX-4
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
tcpO2 (tanscutan) | M1018A | x | x | 1 | x | 1 | 1 | x | x | x | x | x | x | x | x | 1 | 1 | 1 | 1 | Obsolete since 31-Dec-2017 Supported with <= M.0 Not supported with MX500/550
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | x |  | x |  | x |  | x |  | x |  | x |  | x
Philips FAST SpO2 | M1020B #A01 | 2 | x | 2 | x | 2 | 1 | 1/2g) | x | 1/2g) | x | 2 | 2o) | 2 | 2o) | 1/2g) | 1 | 2 | 2o) | MP60/70: For usage in internal Slot B.1 <= required. Two SpO2 channels are supported. M1020B #A04 requires SW ver. >= C.0
Nellcor SpO2 OxiMax compatible | M1020B #A02
|  |  | 2 |  | 2 |  | 2 |  | 2 |  | 2 |  | 2 |  | 2 |  | 2 |  | 2
Masimo SET SpO2 (distributed by Philips) | M1020B #A03/ 867192 #SP3
|  |  | x |  | 1 |  | 1 |  | x |  | 1 |  | x |  | x |  | x |  | x
Nellcor SpO2 module | M1020B #A04
Masimo rainbow SET® SpO2 (distributed by Philips) | 867191 #SP5 | 2p) | x | 2p) | x | 2p) | 1 | 1/2g, p) | x | 1/2g, p) | x | 2p) | 2o) | 2p) | 2o) | 1/2g, p) | 1 | 2p) | 2o) | SW ver. >= L.0 required
|  |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  | 1
|  |  | x |  | 1 |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
EEG | M1027A/B | x | x | 1 | x | 1 | 1 | x | x | 1 | x | 1k) | 1 | 1k) | 1 | 1 | 1 | 1 | 1 | M1027A is not supported with FMX-4 M1027B: aEEG functionality requires SW Rev >= L.0; SW ver. < K.2 supports 2-channel EEG only; SW ver. >= K.2 supports 4-channel EEG
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
Temp | M1029A | x | x | 3 | x | 4 | 4 | 2 | x | 4 | x | 2 | 2 | 4 | 4 | 2 | 2 | 4 | 4
|  |  | x |  | 3 |  | 3 |  | 2 |  | 3 |  | 2 |  | 3 |  | 2 |  | 3
|  |  | x |  | 2 |  | 1 |  | x |  | 2 |  | x |  | x |  | x |  | x
VueLink Auxiliary VueLink Vent VueLink Gas Analyz. VueLink Anesthesia VueLink Auxiliary+ | M1032A #A01 M1032A #A02 M1032A #A03 M1032A #A04 M1032A #A05 | x | x | 2h) | x | 4h) | 4 | x | x | x | x | x | x | x | x | 2 | 2 | 4h) | 4 | Obsolete since 31-Dec-2018 Supported with <= M.0
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 2 |  | 1 |  | x |  | x |  | x |  | x |  | x |  | x
BIS | M1034A/B | xj) | x | 1 | x | 1 | 1 | x | x | 1l) | x | 1l) | 1 | 1l) | 1 | 1 | 1 | 1 | 1 | MP60/70: For usage in internal Slot SW ver. >= E.0 required M1034A is supported with <= M.0 M1034B requires SW ver. >= K.2
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | 1 |  | x |  | 1 |  | x |  | x |  | x |  | x
Recorder | M1116B/C | x | x | 1 | x | 1 | 1 | x | x | 1 | x | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | 1 |  | x |  | 1 |  | x |  | x |  | x |  | x
EC10 IntelliBridge | 865115 | x | x | 2h) | x | 4h) | 4 | 2m) | x | 2m) | x | 4 | 4 | 4 | 4 | 2/3n) | 2/3n) | 4h) | 4 | SW ver. >= G.0 required, the number of modules is driver dependant; Open Interface driver requires SW ver. >= H.15
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 2 |  | 1 |  | x |  | 2 |  | x |  | x |  | x |  | x
ScvO2 | M1011A | x | x | 1 | x | 1 | 1 | x | x | 1 | x | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | SW ver. >= G.0 required ScvO2 and SvO2 are mutually exclusive
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
SvO2 | M1021A | x | x | x | x | 1 | 1 | x | x | x | x | x | x | x | x | 1 | 1 | 1 | 1 | SvO2 and ScvO2 are mutually exclusive Obsolete since 31-Dec-2015 Supported with <= L.0
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
TcG10 | 865298 | x | x | 1 | x | 1 | 1 | 1l) | x | 1 | x | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | SW ver. >= H.0 required
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | 1 |  | x |  | 1 |  | x |  | x |  | x |  | x
Masimo O3 | 867184 | x | x | x | x | x | x | x | x | 1 | x | 1 | 1 | 1 | 1 | x | x | x | x | SW ver. >= N.0 required
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | x |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
Masimo CO2 | 867185 | x | x | x | x | x | x | x | x | 1 | x | 1 | 1 | 1 | 1 | x | x | x | x | SW ver. >= N.0 required
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | x |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
NMT | 865383 | x | x | 1 | x | 1 | 1 | x | x | 1 | x | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | SW ver. >= J.0 required
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
G7m | 866173 | x | x | 1 | x | 1 | 1 | x | x | 1 | x | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | SW ver. >= L.0 required
|  |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x |  | x
|  |  | x |  | 1 |  | x |  | x |  | 1 |  | x |  | x |  | x |  | x
ECG | M1001B | Not compatible with IntelliVue family Only compatible with CMS/V24/26-series
ECG/Resp | M1002B
Pressure | M1006A
Press with analog out | M1006A #C01
NIBP | M1008B
Sidestream CO2 | M1015A
End-Tidal CO2 | M1016A
FIO2 | M1017A
SpO2 Philips FAST | M1020A
BAM | M1022A
Data Transfer | M1235A
For more information regarding compatibility of modules please see SB86200067
a) 3rd pressure capability requires #C30
b) 4th pressure capability requires #C31
c) 3rd pressure capability requires #M05 (not for MX430), 4x pressure capability standard with software N.0 and higher
d) 3rd and 4th pressure capability requires #M06, 4x pressure capability standard with software N.0 and higher
e) MP90 and MX800 >= J.0 support up to 10x pressure capability, all 10x pressures can be M1006B modules with max. 5 per FMS (using two FMS-8)
f) C.O. capability requires #M12 (not for MX430)
g) 2nd SpO2 capability requires #M20, 2x SpO2 capability standard with software N.0 and higher
h) IntelliBridge is the Vuelink replacement. Therefore, either up to 4 Vuelink or up to 4 IntelliBridge can be plugged. Furthermore, IntelliBridge has driver dependent rules.
j) BIS solution for MP20/30 with internal I/O card available
k) M1027B EEG module only
l) M1034B BIS module only
m) IntelliBridge solution for MX400/450/500/550 with internal I/O card available (not for MX430)
n) MX600/700 supports up to 4 modules, even with an M8048A FMS-8. With software >= M.04.00 and option H30 (commercial option H32), 8 modules in total and up to 3 IntelliBridge EC10 modules are supported
o) With two connected FMS only and one module in each FMS
p) Additional Rainbow parameters only available for pulse oximetry label “SpO2”
Module Number | Module Description | X2/MP2, MP5T/5SC | MP5 | MP20/MP30 | MP40/MP50 | MP60/MP70 | MP80/MP90 | MX400 - MX450 | MX500/MX550 | MX600/MX700 | MX800 | MX750 | MX850 | Comments
866173 | G7m | X | X | X | O | O | O | X | O | O | O | O | O | Requires SW Rev >= L.0
M1013A | G1 | X | O | O | O | X | X | O | O | O | O | O | O | Via MIB (RS232); requires SW Rev >= B.1
M1019A | G5 | X | O | O | O | O | O | O | O | O | O | O | O | Via MIB (RS232); requires SW Rev >= D.0
MX750/8501) Software | Active Display (AD75/85)
| N.01.00 | N.01.01 | N.01.04
N.01.00 | O | O | O
N.01.01 | O | O | O
N.01.04 | X | X | O
1) MX850 supports up to two Active Display connections
IntelliVue SW | PIIC iX Revision
| A.0x2) | B.03) | B.013, 4) | B.025) | C.026) | C.036)
<= H.15 | X | X | X | X | X | X
J.01) | O | O | O | O | X | X
J.1 | O | O | O | O | O | O
J.2 | A.02.04 | O | O | O | O | O
K.2 | A.02.13 | O | O | O | O | O
L.04) | A.02.13 | O | O | O | O | O
L.1x | A.02.13 | O | O | O | O | O
M.0x4) | O | O | O | B.02.10 | O | O
N.04) | X | O | O | O | C.02.02 | O
1) Caregroups require monitors on J.1x.xx or higher, see details in SB86201675A
2) PIIC iX A.0x is obsolete since December 2018, see SB86202845A
3) See FCO86201814A
4) Masimo rainbow SET® requires PIC iX B.01 or higher
5) MX750/850 require PIC iX B.02 or higher
6) Encryption between PIC iX and IVPM require monitors on M.04 or higher
Legend:
O | Green boxes show PIIC iX and IVPM system software compatibility and support. The text catalogs are fully compatible and the system supports all the functionality that the monitors and central products provide.
O | Yellow boxes show PIIC iX and IVPM system software compatibility and support. The text catalog on the central station may not be at the latest release. If a customer claims "Unknown" INOPs, a PIIC iX application update to the minimum version indicated in the table above is required, which includes the latest text catalog. Please consult the PIIC iX service documentation (FCO/SB and Service Guide) before implementing any upgrade.
X | Red boxes show PIIC iX and IVPM systems that are not compatible and not supported.
IntelliVue SW | PIIC Revision1)
| G.0 | H.0 | J.0 | K.0 | L.0 | M.0 | N.03) | N.014)
B.0 | O | O | O | X | O | O | O | O
B.1 | O | O | O | X | X | X | X | X
C.0 | O | O | O | O | O | O | O | X
D.0 | O | O | O | O | O | O | O | X
E.0 | O | O | O | O | O | O | O | X
F.0 | O | O | O | O | O | O | O | O
G.0 | O | O | O | O | O | O | O | O
H.15 | X | X | O | O | O | O | O | O
J.0 | X | X | X | X | O | O | O | O
J.1 | X | X | X | X | O | O | O | O
J.2 | X | X | X | X | O | O | O | O
K.2 | X | X | X | X | O | O | O | O
L.x | X | X | X | X | O2) | O2) | O2) | O2)
M.0x | X | X | X | X | X | X | O2) | O2)
N.0 | X | X | X | X | X | X | X | X
1) Masimo rainbow SET® not supported
2) EEG: Only values from unilateral EEG sensor are supported
3) PIIC N.0 is obsolete since March 2016, see SB86202334A
4) PIIC N.01 is obsolete since December 2019, see SB86202990A
Legend:
O | Green boxes show PIIC and IVPM system software compatibility and support. The text catalogs are fully compatible and the system supports all the functionality that the monitors and central products provide.
O | Yellow boxes show PIIC and IVPM system software compatibility and support. The text catalog on the central station may not be at the latest release. If a customer claims "Unknown" INOPs, the text catalog must be upgraded to the latest text catalog release on the central station system before patients are monitored on the central station. Please consult the PIIC Service Guide for further details.
X | Red boxes show PIIC and IVPM systems that are not compatible and not supported.
Note on E296: SW < K.2: # H30 only
Note on F296: SW < K.2: # H30 only
Note on G296: SW < K.2: # H30 only
Note on H296: SW < K.2: # H30 only
Note on I296: SW < K.2: # H30 only
Note on L296: SW < L.0: # H30 only
Note on M296: SW < L.0: # H30 only

## SW CL
Cableless Measurements SW Revisions History
| 865215 | 865216 | 865218 | 866487 | 866488 | 865220 | 865221/22 | Comment
| SpO2 Pod | NBP Pod | Resp Pod | Avalon CL Wide Range Pod | Avalon CL Fetal & Maternal Pod | Charging | Transmitter/ | Related FCO/SB
[6 empty] | Station | Hotspot
A.0 | A.02.06 | A.02.06 | x | x | x | A.02.06 | x | Initial SW
SB86201203A
2010-11-11
A.1 | A.10.68 | A.10.68 | x | x | x | A.10.69 | x | SB86201400A
supports extended battery operating time for Pods
2011-09-15
A.1 Patch | A.10.77 | A.10.77 | x | x | x | - | x | SB86201566A
required when using with bedside
2011-11-30 | [7 empty] | monitors with SW >= Rev. J.0
A.3 | - | - | x | x | x | - | A.30.88 | SB86201703A
865221/22 IntelliVue Cableless Measurement new
2011-12-04 | [7 empty] | Software A.3 enables WLAN functionality
B.02 | B.02.86 | B.02.86 | B.02.86 | x | x | B.02.86 | B.02.86 | SB86201977A
SB86201978A
2013-12-19
B.02 Patch | B.02.90 | B.02.90 | B.02.90 | x | x | - | B.02.90 | FCO86201562A
2014-06-06
B.02 Patch | B.02.94 | - | - | x | x | - | - | SB86202231A
2015-04-21
C.0 | C.01.24 | C.01.24 | C.01.24 | x | x | C.01.24 | C.01.24 | SB86202377A
2016-02-24
D.0 | D.00.93 | D.00.93 | D.00.93 | x | x | D.00.93 | D.00.93 | SB86202601A
D.10.93
2017-06-13
D.0 Patch | - | - | - | x | x | - | D.00.95 | SB86202662A
2017-09-26
D.0 Patch | D.04.16 | D.04.16 | - | x | x | D.04.16 | D.04.16
2018-07-12
D.0 Patch | D.05.23 | D.05.23 | D.05.23 | D.05.23 | D.05.23 | D.05.23 | D.05.23 | SB86202759A
D.15.23
2019-08-01
This overview shows all upgradeable SW revisions released with the CL Measurements family. In the Support Tool some of them are no longer available because they were replaced by newer revisions. Please use always the latest Support Tool to service the devices.
Note on B46: B300-2019-06-25159

## Compatibility CL
CL Compatiblity
CL SpO2/NBP Pod SW | PIIC Revision |  | PIIC iX Revision
| <= L.0 | >= M.0 | >= A.0 via monitor | >= A.0 via PWD/PWM
A.0 | X | O | X | X
A.10.68 | X | O | X | O
A.10.77 | X | O | O | O
B.02 | X | O | O | O
C.0 | X | O | O | O
D.0 | X | O | O | O
CL Respiration Pod SW | PIIC Revision | PIIC iX Revision
| all Revisions | <= A.02, bedside monitor >= J.2 | >= B.0, bedside monitor >= J.2
B.02 | X | O | O
C.0 | X | O | O
D.0 | X | O | O
CL SpO2/NBP Pod SW | TRx4841A 1.4 GHz |  | TRx4851A 2.4 GHz
| Device FW <= C.00.xx | Device FW >= D.00.xx | Device FW <= C.00.xx | Device FW >= D.00.xx
A.0 | X | O | X | O
A.1 | X | O | X | O
B.02 | X | O | X | O
C.0 | X | O | X | O
D.0 | X | O | X | O
Note: The CL Respiration Pod is not compatible with the Telemetry Device TRx4841A/TRx4851A
CL SpO2/NBP Pod SW | MX40
>= A.00.xx
A.0 | X
A.1 | O
B.2 | O
C.0 | O
D.0 | O
CL SpO2/NBP Pod SW | X2/MP21) |  |  |  |  | MP5/MP5SC/MP5T | [6 empty] | Avalon CL
| <= G.0 | H.15 | J.x | >= K.2 | M.0 | <= G.0 | H.x | J.x | >= J.2 | L.1 | M.0 | N.02) | J.3 | L.3
A.0 | X | O | X | X | X | X | O | X | X | X | X | X | X | X
A.10.68 | X | O | X | X | X | X | O | X | X | X | X | X | X | X
A.10.77 | X | O | O | X | X | X | O | O | X | X | X | X | X | X
B.02 | X | O | O | O | X | X | O | O | O | X | X | X | O | X
C.0 | X | O | O | O | X | X | O | O | O | O | X | X | O | X
D.0 | X | O | O | O | O | X | O | O | O | O | O | O | O | O
CL Respiration Pod SW | X2/MP21) |  |  | MP5/MP5SC/MP5T
| <= J.x | >= K.2 | M.0 | < J.2 | >= J.2 | L.1 | M.0 | N.02)
B.02 | X | O | X | X | O | X | X | X
C.0 | X | O | X | X | O | O | X | X
D.0 | X | O | O | X | O | O | O | O
D.1 | X | O3) | O | X | O3) | O3) | O | O
1) X3/MX100 do not support the CL SpO2/NBP/Resp Pod
2) MP5/5SC only
3) Not supported when pediatric patient category is used
CL SpO2/NBP Pod SW | Guardian Software
| <= A.2 | >= B.0 | >= C.0 | >= D.0
A.0 | X | X | X | X
A.1 | O | X | X | X
B.02 | X | O | X | X
C.0 | X | O | O | X
D.0 | X | O | O | O
CL Respiration Pod SW | Guardian Software
| < B.0 | >= B.0 | >= C.0 | >= D.0
B.02 | X | O | X | X
C.0 | X | O | O | X
D.0 | X | O | O | O
D.1 | X | O1) | O1) | O
1) Not supported when pediatric patient category is used
CL Transmitter/Hotspot SW | Guardian Software
| <= A.1 | A.2 | >= B.0 | >= C.0 | >= D.0
A.3 | X | O | X | X | X
B.02 | X | X | O | X | X
C.0 | X | X | O | O | X
D.0 | X | X | O | O | O
CL Respiration Pod | CL Charging Station |  | CL Transmitter/Hotspot
| < B.0 | >= B.0 | < B.0 | >= B.0
>= B.02 | X | O | X | O

## Compatibility XDS
Monitor/XDS Application Compatiblity
NOTE: XDS Application Software update is available free of charge. Printing Solution is integrated in XDS and doesn't require an X-option in the montior
Intellivue Monitor SW2) | Printing Solution | IntelliVue XDS Application
|  | G.0 2/2009 | H.0 | J.0 | J.1 | J.2 | K.2 | L.0 | L.1 | M.x3) | N.0/N.01
B.0 | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL
B.1 | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL
C.0 | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL
D.0 | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL
E.0 | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL | PPCL
F.0 | PAnyType | PAnyType | PAnyType | PAnyType | PAnyType | PAnyType | PAnyType | PAnyType | PAnyType | PAnyType | PAnyType
G.0 | X | XDS | XDS | XDS | XDS | XDS | XDS | XDS | XDS | XDS | XDS
H.0 | X | X | XDS | XDS | XDS | XDS | XDS | XDS | XDS | XDS | XDS
J.0 | X | X | X | XDS | XDS | XDS | XDS | XDS | XDS | XDS | XDS
J.1 | X | X | X | X | XDS | XDS | XDS | XDS | XDS | XDS | XDS
J.2 | X | X | X | X | X | XDS | XDS | XDS | XDS | XDS | XDS
K.2 | X | X | X | X | X | X | XDS1) | XDS1) | XDS1) | XDS1) | XDS1)
L.0 | X | X | X | X | X | X | X | XDS | XDS | XDS | XDS
L.1 | X | X | X | X | X | X | X | X | XDS | XDS | XDS
M.0x | X | X | X | X | X | X | X | X | X | XDS | XDS
N.0 | X | X | X | X | X | X | X | X | X | X | XDS4)
1) See FCO86201610A
2) MP20 - 50 do not support XDS Remote Display connection
3) XDS M.10.00 is not available in FDA regulated territories, see SB86202834A for further details
4) MX750/850 require XDS N.01 (or higher)
X | Not compatible
PPCL | Printing, limited to PCL Printers only
PAnyType | Printing only, any printer type
XDS | Full XDS functionality, includes printing on any type

## Compatibility IGS
Monitor/IGS Compatiblity
Note: For the most recent IGS compatibility information, please visit |  |  |  | IntelliVue GuardianSoftware
IntelliVue MP5/5SC SW | IGS Revision
| A.2 | B.0 | C.0 | C.01 | D.0 | E.0
J.0 | X | X | X | X | X | X
J.1 | O | X | X | X | X | X
J.2 | X | O | O1) | O | O | O
K.2 | X | X | X | X | X | X
L.0 | X | X | X | X | X | X
L.10 | X | X | O | O | O | O
L.12 | X | X | X | O | O | O
M.0x | X | X | X | X | X | X
N.0 | X | X | X | X | X | O
1) IGS C.00.02 required, see SB86202441A

## SW Avalon FM
Avalon FM SW/HW Revisions History
| M2702A/3A | M2704A/5A | Comment
| FM20/30 | FM40/50 | Related FCO/SB
C.0 | C.00.98 |  | Initial SW
(A.00.05/09)
2005-06-01
D.0 | D.00.74 |  | SB86200381
(A.00.05/09)
2006-02-01
D.0 Patch | D.00.75 |  | SB86200534
(A.00.05/09)
2007-01-01
D.0 Patch | D.00.76 |  | SB86200534 (added Hungarian and fix in polish NLS)
| (A.00.05/09) |  | SB86200583
2007-02-01
D.0 Patch | D.00.77 |  | SB86200557 (optimized touch calibration procedure)
| (A.00.05/09) |  | SB86200583
2007-03-01
D.0 Patch | D.00.78 |  | SB86200602
(A.00.05/09)
2007-04-01
F.0 | F.01.51 | F.01.51 | SB86200731
| (A.00.05/09) | (A.00.09) | SB86200732
2008-02-01
F.0 Patch | F.01.52 | F.01.52 | FCO86200878
| (A.00.05/09) | (A.00.09)
2008-04-01
F.0 Patch | F.01.53 | F.01.53 | SB86200950A
| (A.00.05/09) | (A.00.09)
2009-04-01
F.0 Patch | F.01.54 | F.01.54 | FCO86201000A
| (A.00.05/09) | (A.00.09)
2009-05-01
F.0 Patch | F.01.55 | F.01.55 | FCO86201038A FM20/30 only
| (A.00.05/09) | (A.00.09) | FCO8620039A
2009-08-01
F.0 Patch | F.01.56 | F.01.56 | enhanced fetal monitor notes handling when
| (A.00.05/09) | (A.00.09) | used together with OB TraceVue
2010-07-27
G.0 | G.02.12 | G.02.12 | SB86201093A
| (A.00.05/09) | (A.00.09) | FCO86201120A
2010-05-01
G.0 Patch | G.02.13 | G.02.13 | enhanced fetal monitor notes handling when
| (A.00.05/09) | (A.00.09) | used together with OB TraceVue
2010-07-27
G.0 Patch | G.02.14 | G.02.14 | this revision change only enhances internal production
| (A.00.05/09) | (A.00.09) | tests and has no impact on device functionality or
2010-10-20 |  |  | performance of the installed base
G.0 Patch | G.02.19 | G.02.19 | SB86201387A
| (A.00.05/09) | (A.00.09) | FCO86201260A
2011-05-02
G.0 Patch | G.02.21 | G.02.21 | FCO86201355A
| (A.00.05/09) | (A.00.09)
2012-06-19
G.0 Patch | G.02.22 | G.02.22 | SB86202861A
| (A.00.05/09) | (A.00.09)
2018-10-01
J.3 | J.30.58 | J.30.58 | SB86202028A
| (A.00.09/18) | (A.00.09/18)
2014-06-05
J.3 Patch | J.30.59 | J.30.59 | FCO86201584B
| (A.00.09/18) | (A.00.09/18)
2014-08-12
J.3 Patch | J.30.60 | J.30.60 | FCO86201632A
| (A.00.09/18) | (A.00.09/18)
2015-03-11
J.3 Patch | J.30.62 | J.30.62 | SB86202367A
| (A.00.09/18) | (A.00.09/18)
2017-06-13
J.3 Patch | J.30.63 | J.30.63 | SB86202861A
| (A.00.09/18) | (A.00.09/18)
2018-10-01
L.3 | L.33.61 | L.33.61 | SB86202759A
| (A.00.09/18) | (A.00.09/18)
2019-07-25
This overview shows all upgradeable SW revisions released with the Avalon FM. In the Support Tool some of them are no longer available because they were replaced by newer revisions. Please use always the latest Support Tool to service the devices.
Note on B82: B300-2018-09-24193

## FW Avalon FM
Avalon FM FW Revisions History
SpO2
The SpO2 HW is a configurable option of the M2703A. (M2702A has no SpO2 upgrade option)
Option |  | FW Rev | Comment
#B73 |  | A.01.42 | No upgrade for FM 20/30 required (Not affected from the FCO86200673)
|  | A.01.46 | ---
|  | A.01.47 | ---
#A01 FAST SpO2 |  | A.04.12 | for HW B.03.00
A.04.13
A.04.16
NBP
The NBP pump assembly contains a PCB with an upgradeable FW.
Option | FW Rev | Comment
#B71 | A.00.30
#B73 | A.00.31 | SB86200383 M2702A 3A NiBP (M2703-64502) Firmware Revision A.00.31 is released and now available
| A.00.32 | ---
| A.00.33 | ---
| A.00.39 | FCO86201130A: M2702/3/4/5A New NBP Module firmware available. With the NBP Module Firmware A.00.33, a slightly increased variability in the measured systolic value is possible.
| A.00.41 | ---
| A.00.54 | FCO86201795A: M2702/03/04/05A NBP measurement issue in FM20-50 might cause INOPs or aborted NBP measurements - supersedes FCO86201779A
Busmaster
FW Rev | Comment
A.02.00 | Initial revision
A.03.00 | Telemetry Interface support added. FCO86200702 (Upgrade to A.03.02) Despite the implementation of software upgrade A.02.22 (see FCO86200632) a few customer observed that Fetal hear rate (FHR) signal loss is occurring, especially during periods of time when there is notable fetal movement. The monitor does not display FHR numerics nor is there a FHR trace generated by the recorder. However, an audible FHR QRS tone is heard. In some cases the maximum audible signal Volume is not loud enough for some customers.)
A.03.02 | ---
A.04.22 | FM40/50 support FCO86200878: M2702/3/4/5A New Fetal Monitor S/W and Busmaster/API FW are available to fix several issues, 1. When an Avalon CTS is connected to a FM20/30/40/50 with Software Revision F, a sound similar to a fetal heart rate of 240 bpm may be audible if the volume is set to a high value and no real fetal heart rate is being recorded. Neither a heart rate is displayed on the screen nor is a heart rate printed on paper. 2. During internal performance tests it was found that Avalon fetal monitors may print trace data containing small artificial "steps" and also sending such traces to OB TraceVue. This happens under the following conditions: - Avalon FM is connected to an OB TraceVue system and the time of the fetal monitor requires synchronization with the system time and - there is no paper in the fetal monitor but it is sending data to OB TraceVue (bridge paper out functionality) or - the stored data recording functionality is being used to print a patient’s trace onto paper. Realtime recording is not affected by this behavior. Retrofit on failure upgrade to A.04.24
A.04.24 | ---
A.10.04 | ---
A.11.24 | SB86202759A: M2702/3/4/5A New Avalon CL Wide Range/CL Fetal & Maternal pod/Avalon SW Rev. L.3 available
M2734A/5A/6A Transducer
FW Rev | Comment
A.02.24 | Initial revision
A.02.22 | Added dynamic movement suppression to reduce trace spikes
A.02.26 | FCO86200702 Despite the implementation of software upgrade A.02.22 (see FCO86200632) a few customer observed that FHR signal loss is occurring, especially during periods of time when there is notable fetal movement. The monitor does not display FHR numerics nor is there a FHR trace generated by the recorder. However, an audible FHR QRS tone is heard. In some cases the maximum audible signal Volume is not loud enough for some customers.)
A.03.01 | FCO86200879 (Upgrade to A.04.15): M2734/5/6/8A, M2702/3/4/5A, M2720A Improved direct / maternal ECG algorithm is released It was possible that DECG traces were superimposed with artificial spikes of equal length with positive and negative polarity (e.g. above and below the actual trace). Service Recommendation upgrade to A.04.15
A.04.15 | FCO86200879 Service Recommendation Upgrade to A.04.15
A.04.30 | FCO86200974: M2734/5/6/8A, M2702/3/4/5A, M2720A Further improved direct / maternal ECG algorithm is released, supersedes FCO86200879
A.06.31 | Philips has developed modified transducer firmware that will bring the overall monitoring experience for the Avalon series fetal monitors closer to that of the Series 50 fetal monitors. All countries except US see FCO86201118A: M2731-60001/M2732-60001/M2736A new ultrasound Firmware available. FCO86201218A also for US customers
A.06.45 | FCO86201835A: M2734B/M2735A - Toco Equip Malf - new Transducer Firmware available
867245/46/47/48/49 Transducer
FW Rev | Comment
K.01.08 | Initial revision SB86202936A: 867245/6/7/8/9 New wired Avalon transducers and transducer cable kits available
M2738A Patient Module
FW Rev | Comment
A.01.15 | Initial revision
A.01.17 | FCO86200702: M2702/3A, M2734A/35A/36A/38A, M2731-60001 New Busmaster and Transducer Firmware available
A.01.18 | FCO86200879 (Upgrade to A.04.15): M2734/5/6/8A, M2702/3/4/5A, M2720A Improved direct / maternal ECG algorithm is released It was possible that DECG traces were superimposed with artificial spikes of equal length with positive and negative polarity (e.g. above and below the actual trace).. Service Recommendation upgrade to A.01.18
A.01.24 | FCO86200974: M2734/5/6/8A, M2702/3/4/5A, M2720A Further improved direct / maternal ECG algorithm is released, supersedes FCO86200879 Service Recommendation Upgrade to A.01.24
M2731-60001 / M2732-60001 CTS Connector Cable
FW Rev | Comment
A.01.16 | Initial revision
A.01.29 | FCO86200702 (Upgrade to A.01.29) Despite the implementation of software upgrade A.02.22 (see FCO86200632) a few customer observed that FHR signal loss is occurring, especially during periods of time when there is notable fetal movement. The monitor does not display FHR numerics nor is there a FHR trace generated by the recorder. However, an audible FHR QRS tone is heard. In some cases the maximum audible signal Volume is not loud enough for some customers.)
A.01.32 | FCO86200731 (Upgrade to A.01.32) In rare cases when the Avalon CTS (M2720) and the FM20/30 (M2702/3A) are powered up synchronously, e.g. if both instruments are installed on a cart with a power master switch, the Avalon CTS interface cable (M2731-60001) did not work initially.
A.01.38 | FCO86200879 (Upgrade to A.04.15): M2734/5/6/8A, M2702/3/4/5A, M2720A Improved direct / maternal ECG algorithm is released It was possible that DECG traces were superimposed with artificial spikes of equal length with positive and negative polarity (e.g. above and below the actual trace). Service Recommendation upgrade to A.01.38
A.01.44 | FCO86200974: M2734/5/6/8A, M2702/3/4/5A, M2720A Further improved direct / maternal ECG algorithm is released, supersedes FCO86200879 Service Recommendation Upgrade to A.01.44
A.01.46 | All countries except US see FCO86201118A: M2731-60001/M2732-60001/M2736A new ultrasound Firmware available
A.01.50 | Worldwide, with the following exceptions: Islamic Republic of Iran, Democratic People's Republic of Korea, Myanmar and USA *) see FCO86201175A; FCO86201218A also for US customers
866074 Avalon CL Basestation
Connector Cable FW Rev |  | Comment
A.10.06 |  | Initial revision (including OB CL Transducer FW A.10.02)
A.10.07 |  | (including OB CL Transducer FW A.10.03)
A.10.10 |  | (including OB CL Transducer FW A.10.10)
B.01.07 |  | (including OB CL Transducer FW B.01.07) SB86202759A: M2702/3/4/5A New Avalon CL Wide Range/CL Fetal & Maternal pod/Avalon SW Rev. L.3 available
SRR
FW Rev | Related Products |  | Comment
B.02.60 | 866060/62/64/66/74 |  | Initial revision
D.00.42 |  |  | SB86202759A: M2702/3/4/5A New Avalon CL Wide Range/CL Fetal & Maternal pod/Avalon SW Rev. L.3 available

## Revision History
Revision History
Revision | Description | Date | Initiator
SB86202000K | Added new patches for IntelliVue L.0, M.04 and N.0 | 2021-02-23 | C.S.
SB86202000J | IntelliVue N.0/N.01 SW, Avalon L.3 SW, new MX750/850, AD75/85, and FMX-4, added patches for IntelliVue M.04 and Avalon G.0/J.3 | 2019-10-29 | C.S.
SB86202000H | IntelliVue M.04 SW, added patches for G.0, H.15, J.0, J.1, K.2, L.0, L.12, and M.0 | 2018-09-07 | C.S.
SB86202000G | Obsolete products removed, added patches for J.x, K.2, L.0, L.12, and M.0 | 2017-12-04 | C.S.
SB86202000F | IntelliVue M.0 SW, CL devices D.0 SW, IGS C.01, added patches for J.0, J.1, K.2, L.0 and L.12 | 2017-06-19 | C.S.
SB86202000E | IntelliVue L.1 SW, CL devices C.0 SW, added patches for K.2 and L.0 | 2016-11-25 | C.S.
SB86202000D (L.0.1) | IntelliVue L.0 SW, added patches for J.2, J.3 and K.2 | 2015-08-04 | C.S.
SB86202000C (K.0.5) | Added new patches for J.1, J.2, K.1, K.2 and J.22 Epidemic Triage protocol, IntelliVue Guardian Software compatibility added | 2014-11-24 | C.S.
SB86202000B (K.0.4) | IntelliVue K.2 SW, added new patches for J.2/J.3 | 2014-09-25 | C.S.
SB86202000A (K.0.3) | IntelliVue K.1 and J.2 SW, CL devices B.02 SW, Avalon J.3 SW, new MX500/550, added new patches for J.0, J.1, K.1, XDS J.2, from now on Matrix is distributed in Service Bulletin format | 2014-06-30 | C.S.
K.0.2 | Added new patches for J.1 and K.0 | 2013-11-28 | C.S.
K.0.1 | Introducing K.0 SW, new MX400/450 | 2013-09-10 | C.S.
J.1.1 | Added new patch for H.15/J.0/J.1, MP2/X2 HW Rev. B, CL A.3 | 2013-06-07 | C.S.
J.1.0 | Added new SW J.1, Nellcor SpO2 module | 2012-11-15 | C.S.
J.0.2 | Added new Patch SW for J.0 and H.15, new NMT module, CL SW history | 2012-08-10 | C.S.
J.0.1 | Introducing J.0 SW, TcG10 and CL included | 2012-06-27 | C.S.
H.0.7 | Added new Patch SW for H.15/IIC N.0 compatibility, added M3015B | 2011-10-27 | C.S.
H.0.6 | Added new Patch SW, MX600/700/FMS4 / Modules and MX600/700 and SRR FW dependency to H.0 |  | A.M / I.H
H.0.5 | Added 12ld upgrade, new WLAN FW in FW section, added new Patch SW |  | A.M
H.0.4 | Added SW Patch for ECG Application |  | A.M
H.0.3 | Added SW Patch and corrected number of Temp modules in Internal Slot of MP70 |  | A.M
H.0.2 | M8001A/2A Main baord revision changed for H.0 compatibility |  | A.M
H.0.1 | Introducing H.0 SW / Not supported SW revisions have been erased |  | A.M
G.0.9 | New SW and new SpO2 FW |  | A.M
G.0.8 | Revised Compatibility Matrix for Modules |  | A.M / I.H.
G.0.7 | Added new SW Patch. Comment on MMS/M3 compatibility and Intellibridge. PIC Rev L is now compatible with IV Rev B.0. NLS Revision. XDS compatiblity added. FW for WLAN Adapter for X2/MP2 |  | A.M
G.0.6 | Added new Fetal SW |  | A.M
G.0.5 | EEG not supported in internal slot of MP60/70 |  | A.M
G.0.4 | New SW for MMS and B.0 Patch for ELO touch and correctedFetal SW |  | A.M
G.0.3 | New SW for Hungarian language and new Fetal SW |  | A.M
G.0.2 | Add new G.0 Patch release for IV and Fetal |  | A.M
G.0.1 | Added comment in ECG FW F.0 and higher |  | A.M
G.0 | G.0 Release and modifications in Compatibility |  | A.M
0.5 | Added new F.0 Patch |  | A.M
0.4 | Added B.0 requirement for MP60/70 for Vuelink |  | A.M
0.3 | Modules compatibility change + Fetal monitoring included |  | A.M
0.2 | IV SW B.1 not compatible with IC rev K. corrected |  | A.M
0.1 | Initial Release |  | A.M