**TC 50 Setup guide**

* Create a user name for the device (in accordance with IT Wi-Fi LOGIN naming convention) E.g.: JHHWARDF355106 and get IT to create an account.
* Power unit & insert USB stick.

Tap on the following buttons:

“Setup” → “Configure Cardiograph Default Settings” → “Save/Load Settings” → “Load”.

Tap “Browse” and find the XML file, tap “OK”.

Enable “√” in “load network settings from” box & browse and find DAT file.

Tap “OK”

Tap “OK”

Unit will load files & reboot.

* After the unit reboots:

“Setup” → enter password (biomed) → “Configure Cardiograph Default Settings → “Institution”.

Enter the BME number in “Device ID”

Delete (DEL) Institution name & enter site name & tap “ADD” (e.g.: JHH ED 01)

Tap “Default”

Tap “Exit”

Tap “Yes”

* “Configure ECG Network Settings” → “Wireless Network” → “Configure” → “Admin Login” → enter password (SUMMIT) → “OK” → “Profile”.

The profile should read HNE & the SSID, Encryption & EAP type fields should be filled. (SSID: ND0QpwJP7usK8e0C0JzW3wqCZruZ0W8j )

(PASSKEY: jZDJ9rN6). “Credentials” enter I.T. Login from spreadsheet under User Name.

* Verify timeserver (HNEJHHDC02) and daylight saving settings are correct.
* While still in setup touch time and verify its set to Sydney time. Perform manual time sync.



